## DataBase
Manufacturer | User - Buyer | Project | reference | year_ref | link_reference | Model | Commercial | living | Manufacturer Web link | type(opc, nephelometer, electrochemical, metal oxide) | OEM*/ Sensor system1 | OEMFirst provider | Pollutant | Sensor result unit | Availability of raw data | Open source correction or black box? | Field / Lab | site | Response Time (s) | Time resolution (s) | Averaging time | units (avg_time) | R2_calib | intercept(a0) | a1(slope) | a2 | a3 | a4 | a5 | model type calibration | calibration type equation2 | Int_O3 | Int_T | Int_RH | Int_NO | Int_NO2 | Int_CO | Int_CO2 | Int_day | X | Y | model type regression (comparison) | r2 | intercept | slope | RMSE | MAE | BIAS | nRMSE (%) | min data recovery(%) | max data recovery(%) | Intra-model variability mean Field Value (µg/m3) | T (°C) | RH (%) | % failure | NO2, Sensitivity Coeff (ppb/ppb or ppm) | NO, Sensitivity Coeff (ppb/ppb or ppm) | CO, Sensitivity Coeff (ppb/ppm) | O3, Sensitivity Coeff (ppb/ppb or ppm) | SO2, Sensitivity Coeff (ppb/ppb or ppm) | Temp, Sensitivity Coeff (ppb/°C) | RH, Sensitivity Coeff (ppb/%) | U | type of U | RSD (100-1000 µg/m3) (%) | RSD (< 100 µg/m3) (%) | min Accuracy | max Accuracy | units Accuracy | LoD | units (LoD) | LoD (100-1000 µg/m3) µg/m3 | LoD (< 100 µg/m3) µg/m3 | RANGE | units RANGE | precison | units (precision) | repteability (at 100 ppb) | Reproducibility between sensors (R2) | Dss-drift | Dls-drift (ppb/day) | Dep. Part. Composition | impact Part. Compositions (%) | Dep. Part. Size | impact part. Size (%) | Dep. RH | impact RH (%) | Dep. T | impact T (%) | Long Term Performance | Price | currency | link_reference | comments | operator (added by)
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8669 | 4.1665 | 0.9373 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8412 | 4.9107 | 0.9263 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.7842 | 3.5351 | 0.9539 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8481 | 4.0137 | 0.9872 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8745 | 4.1148 | 0.9418 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8484 | 4.815 | 0.9318 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8042 | 3.4997 | 0.9622 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8562 | 3.9516 | 0.9928 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9225 | 3.6646 | 0.9773 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9126 | 4.819 | 0.9559 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8645 | 3.6032 | 0.9468 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9233 | 3.5683 | 1.0555 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8435 | 4.0564 | 0.9077 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8302 | 3.7918 | 0.9194 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.533 | 3.4347 | 1.3074 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8372 | 4.0173 | 0.9689 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.9247 | 3.9097 | 0.9111 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.9318 | 3.8435 | 0.9201 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.5754 | 3.5293 | 1.1949 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8849 | 4.4864 | 0.8085 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7477 | -1.7229 | 1.1845 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7831 | -0.2971 | 0.9549 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.9153 | 4.5088 | 0.9153 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.9801 | 2.0031 | 0.9801 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7787 | -2.2543 | 1.2225 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8063 | 0.8664 | 0.978 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.2058 | 2.6131 | 1.1443 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8144 | 1.9001 | 0.9913 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8313 | -1.4757 | 1.1741 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8155 | -3.8834 | 1.0972 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.3971 | -1.1867 | 1.5499 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | unknow sensor | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8324 | 3.7397 | 0.8904 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.823 | 3.5456 | 0.9781 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8365 | 3.2115 | 1.0069 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8351 | 3.5039 | 0.983 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8488 | 3.1674 | 1.0119 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9282 | 3.44 | 0.9876 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9469 | 3.1393 | 1.0138 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3126 | 15.781 | 1.1348 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0682 | 21.925 | 0.8345 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3461 | 15.784 | 1.135 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3713 | 15.106 | 1.2003 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0659 | 22.221 | 0.7287 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4418 | 15.703 | 1.1281 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4776 | 15.048 | 1.1921 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.705 | 5.217 | 5.681 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.818 | 3.59 | 5.823 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.634 | 4.037 | 1.34 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.659 | -2.59 | 8.88 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.794 | -5.29 | 9.182 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.654 | 6.11 | 0.713 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.451 | -25.326 | 16.939 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | 0.567 | -34.899 | 18.658 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.489 | 10.873 | 0.48 |  |  |  |  | 100 | 100 | 39 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.706 | 5.179 | 5.683 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.822 | 3.535 | 53838 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.672 | 3.745 | 1.419 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.663 | -2.7402 | 8.934 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.805 | -5.472 | 9.2455 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.698 | 5.652 | 0.7534 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.465 | -25.426 | 16.915 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.597 | -35.437 | 18.763 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [8 empty] | PM=a1*Rs+a0 | [8 empty] | Sensor | GRIMM | linear | 0.524 | 9.909 | 0.505 |  |  |  |  | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.545 | 4.988 | 5.162 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.808 | 2.754 | 5.951 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.583 | 3.664 | 1.451 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.644 | -4.1935 | 9.341 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.756 | 4.809 | 0.83 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.883 | 3.89 | 1.0306 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.579 | -27.738 | 17.412 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.853 | -43.518 | 20.86 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.7664 | -42.362 | 20.108 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.673 | 3.319 | 0.86 |  |  |  |  | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.38 | 8.882 | 0.665 |  |  |  |  | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.395 | 8.804 | 0.594 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.4175 | 4.39 | 0.647 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.527 | 3.647 | 0.711 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 1 | hour | 0.434 | 4.092 | 0.616 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.904 | 2.617 | 0.906 |  |  |  |  | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.826 | 5.982 | 0.952 |  |  |  |  | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.763 | 6.5 | 0.796 |  |  |  |  | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.658 | 3.343 | 0.7368 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.919 | 1.874 | 0.857 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | FIELD | rural |  | 15 | 24 | hour | 0.822 | 1.979 | 0.8245 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | open source | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -7.8 | 0.64 |  |  |  |  | 42 | 100 |  | 20 | 40 | [12 empty] | 0.3 | 43.8 | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -8.98 | 0.56 |  |  |  |  | 42.4 | 100 |  | 20 | 40 | [12 empty] | -33.2 | 1780% | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | open source | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -1.72 | 0.97 |  |  |  |  | 9.5 | 96.8 |  | 20 | 40 | [12 empty] | 85.1 | 9590% | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.878 | -0.96 | 0.906 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.85 | -0.56 | 0.683 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.421 | 5.768 | 0.365 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.806 | 3.075 | 0.599 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.782 | 3.49 | 0.489 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.591 | 5.854 | 0.375 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.156 | 26.86 | 0.506 |  |  |  |  | 66 | 86 | 30 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.198 | 23.97 | 0.46 |  |  |  |  | 66 | 86 | 39 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.227 | 22.8 | 0.457 |  |  |  |  | 66 | 86 | 38 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | -0.8463 | 0.896 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | -0.6169 | 0.685 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6896 | 1.6093 | 0.592 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.821 | 2.948 | 0.599 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.805 | 3.479 | 0.488 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.773 | 3.12 | 0.483 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.94 | -1.07 | 0.89 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.93 | -1.049 | 0.694 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.961 | -0.9156 | 0.719 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | 2.12 | 0.613 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.867 | 2.57 | 0.507 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.933 | 1.843 | 0.526 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.812 | 4.36 | 0.6206 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.802 | 4.904 | 0.506 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.731 | 5.303 | 0.501 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.88 | 4.06 | 0.629 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.86 | 4.351 | 0.523 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.89 | 3.75 | 0.55 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.258 | 33.69 | 0.639 |  |  |  |  | 66 | 86 | 30 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.258 | 32.435 | 0.549 |  |  |  |  | 66 | 86 | 39 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.279 | 31.546 | 0.604 |  |  |  |  | 66 | 86 | 38 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.868 | -0.872 | 0.877 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.866 | -0.789 | 0.888 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8661 | -1.2827 | 0.842 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.78 | 3.257 | 0.602 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7758 | 3.385 | 0.621 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7803 | 2.88 | 0.589 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.113 | 27.65 | 0.546 |  |  |  |  | 87 |  | 26 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.131 | 26.616 | 0.585 |  |  |  |  | 87 |  | 26 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.137 | 23.503 | 0.533 |  |  |  |  | 87 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.888 | -0.8897 | 0.876 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8857 | -0.8077 | 0.887 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8838 | -1.2893 | 0.8404 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.798 | 3.27 | 0.6015 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7961 | 3.44 | 0.617 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.792 | 2.9267 | 0.588 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.936 | -1.289 | 0.8795 |  |  |  |  | 82 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.936 | -1.306 | 0.8947 |  |  |  |  | 82 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.932 | -1.8971 | 0.8538 |  |  |  |  | 82 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.905 | 2.206 | 0.6217 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.901 | 2.3672 | 0.6391 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9021 | 1.8564 | 0.3074 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.83 | 4.091 | 0.638 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.827 | 4.233 | 0.657 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.825 | 3.731 | 0.623 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.912 | 3.808 | 0.645 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9122 | 3.9416 | 0.6652 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9106 | 3.426 | 0.6315 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0967 | 40365 | 0.593 |  |  |  |  | 87 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0952 | 40.97 | 0.605 |  |  |  |  | 87 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  | Portable | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0933 | 40.477 | 0.557 |  |  |  |  | 87 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AirBase |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CanarIT | Y | N |  | nephelometer | sensor system |  | PM | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0394 | 6.5448 | -0.1023 | [45 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
AirBase |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CanarIT | Y | N |  | nephelometer | sensor system |  | PM | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0044 | 6.4979 | -0.1013 | [45 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0576 | 13 | 0.76 | 37 |  |  |  | 93 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.16 | 15 | 0.74 |  |  |  |  | 96 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.1225 | 10 | 0.62 |  |  |  |  | 61 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.76 | 1.76 | 0.22 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.53 | 2.71 | 0.21 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.81 | 0.32 | 0.23 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.67 | 4.62 | 0.16 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | open source | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.61 | 2.47 | 0.13 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | 3.9 | 0.2 | [19 empty] | 6.6 | 6.6 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.94 | 0.7 | 0.03 | [19 empty] | 9.5 | 9.5 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | 5 | 1.7 | [19 empty] | 15 | 15 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | -15 | 0.5 | [19 empty] | 4.8 | 4.8 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.94 | -19 | 0.04 | [19 empty] | 8.7 | 8.7 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | 24 | 0.9 | [19 empty] | 14 | 14 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | -18 | 0.5 | [19 empty] | 4.2 | 4.2 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.96 | -17 | 0.05 | [19 empty] | 6.1 | 6.1 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | LAB |  |  |  | 3 | min | [17 empty] | SMPS | Sensor | linear | 0.99 | 105 | 1.6 | [19 empty] | 16 | 16 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.2025 | -0.3 | 0.4 | 1.08 |  |  |  | 76 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.115 | -1.66 | 0.49 |  |  |  |  | 71 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0121 | 0.6 | 0.07 |  |  |  |  | 82 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.2209 | 2.98 | 0.45 | 1.01 |  |  |  | 76 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.4624 | -1.06 | 0.54 |  |  |  |  | 71 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.04 | 2.86 | 0.12 |  |  |  |  | 82 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.86 | 0.02 | 3.19 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.91 | 0.01 | 2.63 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.63 | 0.04 | 3.53 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.76 | 0.03 | 3.02 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.57 | 0.03 | 2.29 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor |  | 0.67 | 0.02 | 2.06 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [8 empty] | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor |  |  |  | 1.12 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [8 empty] | non linear model including RH |  |  | RH | [5 empty] | GRIMM | Sensor |  |  |  | 1.3 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [8 empty] | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor |  |  |  | 0.905 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | open source | FIELD | urban |  | 15 | 1 | hour | [8 empty] | non linear model including RH |  |  | RH | [5 empty] | GRIMM | Sensor |  |  |  | 1.295 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.486 | 1.955 | 0.0015 |  |  |  |  | 100 |  | 6100 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.436 | 1.8609 | 0.0034 |  |  |  |  | 100 |  | 2900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.502 | 1.69 | 0.0025 |  |  |  |  | 100 |  | 3900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3953 | 6.71 | 0.0364 |  |  |  |  | 100 |  | 530 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5017 | 4.1087 | 0.093 |  |  |  |  | 100 |  | 210 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.147 | 23.896 | 0.0429 |  |  |  |  | 100 |  | 50 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.628 | 1.4255 | 0.0016 |  |  |  |  | 100 |  | 6100 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.54 | 1.1799 | 0.0037 |  |  |  |  | 100 |  | 2900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.605 | 1.3708 | 0.0026 |  |  |  |  | 100 |  | 3900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.43 | 9.85 | 0.0303 |  |  |  |  | 100 |  | 530 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.588 | 6.245 | 0.0839 |  |  |  |  | 100 |  | 210 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.3939 | 21.212 | 0.0867 |  |  |  |  | 100 |  | 50 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | OPC | sensor system | NuWave Sensors - Cair | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.15 |  |  |  | 125.1 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Airviz |  | Carnegie Mellon | US-EPA_2015 | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.039 | 2.76 | 0.146 | [45 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | temperature related correction | JRC-FK
Airviz |  | Carnegie Mellon | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts | N | black box | FIELD | rural |  | 1 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0342 | 119.11 | -6.5434 | [8 empty] | > 95% (spike) | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Airviz |  | Carnegie Mellon | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 1e-05 | 6.7895 | 0.0599 | [8 empty] | > 95% (spike) | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.732 | 2.885 | 0.726 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7579 | 1.99 | 0.8029 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7384 | 2.9473 | 0.7151 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.846 | 2.91 | 0.711 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8705 | 1.976 | 0.791 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | nephelometer | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.852 | 2.905 | 0.7066 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos - DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | 0.814 | 3.977 | -8e-12 | 5e-05 |  |  |  | quad | PM=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | ` |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM | Dylos - DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM |  | 0.83 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM | Dylos - DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 1 | hour | 0.632 | 4.1744 | -1e-11 | 4e-05 |  |  |  | quad | PM=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | BAM | [8 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM | Dylos - DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM |  | 0.89 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM | Dylos - DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM |  | 0.81 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/dylos---lab-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM | Dylos - DC1100 PRO | PM2.5-0.5 | #/L | Y | open source | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.89 | -23404 | 0.93 |  |  |  |  | 97 | 100 |  | 20 | 40 | [12 empty] | -4147 | -117 | [22 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/dylos---lab-evaluation.pdf?sfvrsn=2 |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 | Y | Y |  | OPC | sensor system | Dylos DC1100 | PM2.5-0.5 | pt 0.01cf−1 | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.33 | 8.08 | 0.0086 | [6 empty] | 11.71 | [12 empty] | 0.83 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 PRO | Y | Y |  | OPC | sensor system | Dylos - DC1100 PRO | PM2.5-0.5 | pt 0.01cf−1 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.6 | 14.8 | 0.00164 | -0.0751 | -0.035 |  |  | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day |  |  | RH |  |  |  |  | day | Sensor | BAM | linear | 0.45 | 7.2 | 0.0015 | [6 empty] | 11.67 | [12 empty] | 0.77 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-3 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 PRO | Y | Y |  | OPC | sensor system | Dylos - DC1100 PRO | PM2.5-0.5 | pt 0.01cf−1 | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.4 | 7.78 | 0.0013 | [6 empty] | 11.64 | [12 empty] | 0.85 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.94 | 2.56 | 888.26 | [6 empty] | 18.3 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.94 | 2.58 | 811.52 | [6 empty] | 17.6 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-4 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.82 | 2.91 | 955.26 | [6 empty] | 18.2 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | OEM | Envea - CairClip O3/NO2 | O3 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.709 | 26.6 | -0.175 |  |  |  |  | MLR | O3 = a0 + a1*Rs - a2*RH |  |  | RH | [5 empty] | Sensor | REF | linear | 0.88 | 12.451 | 0.88 | [6 empty] | 18.6 | [35 empty] | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-4 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | OEM | Envea - CairClip O3/NO2 | O3 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.68 | 12.203 | 0.68 | [6 empty] | 16.7 | [35 empty] | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
Dylos |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system |  | PM2.5-0.5 | counts | N | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.533 | 86531 | 21562 | [8 empty] | > 95% (increase) | [31 empty] | YES |  | NO |  |  | 300 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Dylos |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system |  | PM2.5-0.5 | counts | N | open source | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.5483 | 51574 | 21368 | [8 empty] | > 95% (increase) | [31 empty] | YES |  | NO |  |  | 300 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7396 | -152 | 64 | 0.15 |  |  |  | 82 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.6084 | -1182 | 428 |  |  |  |  | 72 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5329 | -941 | 431 |  |  |  |  | 77 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.16 | 5.5 | 1.3 | 0.1 |  |  |  | 82 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.1089 | 73 | 5.7 |  |  |  |  | 72 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0729 | 84 | 4.9 |  |  |  |  | 77 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.88 | 67 | 0.95 |  |  |  | 13.4 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | 0.96 | 30.1 | 0.0056 | 1.99 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | [49 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.67 | 77.65 | 0.7 |  |  |  | 13.4 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.87 | 73 | 0.89 |  |  |  | 13.4 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | 0.93 | 38.86 | 0.0052 | 1.86 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | [49 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.7 | 82.21 | 0.68 |  |  |  | 13.4 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.94 | 0.91 | 0.1 |  |  |  | 46.1 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | 0.82 | 0.42 | -0.000154 | 0.15 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | linear | [48 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.89 | 2.1 | 0.094 |  |  |  | 46.1 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.93 | 0.99 | 0.11 |  |  |  | 46.1 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | 0.78 | 0.42 | -0.000184 | 0.17 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | linear | [48 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.85 | 2.27 | 0.1 |  |  |  | 46.1 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | #/cm3 | Y | open source | LAB |  |  |  | 1 | min | 0.98 | 98 | 8.6 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | SMPS | Sensor | linear | 0.98 | 10 | 1.2 |  |  | -18 |  |  |  | 3.1 | [13 empty] | 8 | [15 empty] | YES | 350 | YES | [6 empty] | 475 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #/m3 | Y | open source | LAB |  |  |  | 1 | min | 0.98 | -43.6 | 0.018 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | DustTrak | linear | [30 empty] | 1 | [17 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Smoke and PLS | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #/m3 | Y | open source | FIELD | traffic |  |  | 24 | hour | [17 empty] | Sensor | DustTrak | linear | 0.99 | 0.008 | 0.51 | [45 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM3 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.98 |  |  | 0.8 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.99 |  |  | 0.6 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.78 |  |  | 4 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.58 |  |  | 3.5 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Steinle et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #ft-3 (PCN) | Y | open source | FIELD | urban |  |  | 1 | hour | 0.9 | 1.29 | 1.11e-05 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | TEOM | [10 empty] | 6.5 | [29 empty] | NO | [8 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub |  | JRC-FK
Dylos |  |  | Steinle et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #ft-3 (PCN) | Y | open source | FIELD | urban |  |  | 1 | hour | 0.7 | 4.75 | 2.8e-05 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | TEOM-FDMS | [10 empty] | 6.5 | [29 empty] | NO | [8 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub |  | JRC-FK
Dylos |  |  | Han et al,.2016 | 2016 | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #ft-3 (PCN) | Y | open source | FIELD | urban |  |  | 1 | min | 0.778 | -2.8197 | 0.0071 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | GRIMM | [10 empty] | 11.3 | [31 empty] | YES | 300 | YES | 110 |  |  |  | 475 | USD | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true |  | JRC-FK
Dylos |  |  | Han et al,.2016 | 2016 | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | #ft-3 (PCN) | Y | open source | FIELD | urban |  |  | 1 | min | 0.481 | -1.0689 | 0.0544 |  |  |  |  | linear | PM10-2.5=a0+a1*Rs | [8 empty] | Sensor | GRIMM | [10 empty] | 4.8 | [31 empty] | YES | 400 | YES | 150 |  |  |  | 475 | USD | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #ft-3 (PCN) | Y | open source | LAB |  |  |  | 1 | min | 0.94 | 0.4922 | 0.7895 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | #ft-3 (PCN) | Y | open source | LAB |  |  |  | 1 | min | 0.77 | 0.01422 | 0.9114 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | #ft-3 (PCN) | Y | open source | LAB |  |  |  | 1 | min | 0.439 | -4.053 | 3.417 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10-2.5 | #ft-3 (PCN) | Y | open source | LAB |  |  |  | 1 | min | 0.00499 | 0.6443 | -0.001876 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour |  | 0.55 | 6.926 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | GRIMM | linear | 0.84 | [47 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour |  | 1.603 | 3.357 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | GRIMM | linear | 0.73 | [47 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Dacunto et al. 2015 | 2015 | https://www.ncbi.nlm.nih.gov/pubmed/26487426 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5 | #/0.01ft3 | Y | open source | LAB |  |  |  | 1 | min | 0.9 |  | 3.91e-06 | 1.803 |  |  |  | exp | PM2.5=a1*Rs² | [59 empty] | 475 | USD | https://www.ncbi.nlm.nih.gov/pubmed/26487426 |  | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | OEM |  | PM2.5-0.5 | mg/m3 | N | open source | LAB |  |  |  | 24 | hour | [7 empty] | linear | [9 empty] | Sensor | DustTrack |  | 0.99 | 0.008 | 0.51 | [26 empty] | < 1 | [18 empty] | 475 | USD | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract |  | JRC-FK
Envea |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CairClip PM2.5 | Y | Y | https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance | nephelometer | OEM | CairPol | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0117 | 0.1311 | -0.0053 | [7 empty] | > 19.8 | > 95% (spike) | [36 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Envea |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CairClip PM2.5 | Y | Y | https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance | nephelometer | OEM | CairPol | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0637 | 2.6107 | -0.2289 | [7 empty] | > 19.8 | > 95% (spike) | [36 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=6 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.573 | 6.087 | 0.43 |  |  |  |  | 96.3 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=7 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.546 | 5.71 | 0.512 |  |  |  |  | 96.3 |  | 18 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=7 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=8 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.547 | 4.514 | 0.42 |  |  |  |  | 96.3 |  | 22 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=10 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.657 | 6.511 | 0.414 |  |  |  |  | 96.3 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.027 | 4.5149 | 0.0352 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1057 | 2.864 | 0.0829 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0189 | 4.654 | 0.0284 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0985 | 8.3815 | 0.1097 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.2159 | 5.2826 | 0.1931 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.086 | 8.522 | 0.0988 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0339 | 87205 | 806.93 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1224 | 51778 | 1823 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0246 | 90080 | 662 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0343 | 87660 | 813.64 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1232 | 52084 | 1833 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0249 | 90539 | 668 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=11 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.583 | 6.58 | 0.462 |  |  |  |  | 96.3 |  | 18 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=11 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=12 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.568 | 6.07 | 0.359 |  |  |  |  | 96.3 |  | 22 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=9 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.65 | 9.318 | 0.138 |  |  |  |  | 96.3 | 100 |  | 20 | 40 | [12 empty] | -1086-(-408) over 10-50 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=9 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.66 | 9.021 | 0.18 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.66 | 9.7253 | 0.1038 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.685 | 9.2526 | 0.1405 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.695 | 8.94 | 0.183 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6918 | 9.659 | 0.1053 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.737 | 8.819 | 0.15 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.744 | 8.5842 | 0.1956 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7415 | 9.225 | 0.113 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7365 | 82050 | 7.907 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.74 | 77809 | 9.43 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7458 | 89100 | 6.55 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.758 | 80754 | 8 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.763 | 76425 | 9.55 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.769 | 87820 | 6.6426 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.89 | 61779 | 9.379 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.893 | 57023 | 11.165 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9 | 70633 | 7.74 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.659 | 10.359 | 0.175 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0022 | 9.2911 | 0.0357 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0166 | 7.8293 | 0.0982 |  |  |  |  | 100 |  | 20 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0203 | 7.752 | 0.1013 |  |  |  |  | 100 |  | 20 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.668 | 9.707 | 0.2317 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6622 | 10.908 | 0.1307 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.767 | 9.499 | 0.1985 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.77 | 8.988 | 0.2599 |  |  |  |  | 99 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7684 | 10.098 | 0.148 |  |  |  |  | 99 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.87 | -8.52 | 0.16 |  |  |  |  | 100 |  |  | 20 | 40 | [36 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/L | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.77 | -3000000 | 5.66 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -1086 | -408 | % |  |  | 0 | [18 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.42 | 5.24 | 0.002 | [6 empty] | 11.2 | [12 empty] | 0.62 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.5 | 12.2 | 0.00203 | -0.0693 | -0.016 |  |  | MLR | NO2=a0+a1*Rs+a2*T+a3*RH |  | T | RH | [5 empty] | Sensor | BAM | linear | 0.43 | 6.01 | 0.0018 | [6 empty] | 11.48 | [12 empty] | 0.58 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-3 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.43 | 6.4 | 0.0017 | [6 empty] | 11.45 | [12 empty] | 0.61 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.25 | 5.52 | 0.06 |  |  |  |  | 88.7 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.21 | 5.98 | 0.05 |  |  |  |  | 95.4 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.33 | 4.15 | 0.06 |  |  |  |  | 92.4 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.67 | -323 | 273 | 0.06 |  |  |  | 68 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7056 | -1234 | 278 |  |  |  |  | 68 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.67 | -352 | 322 |  |  |  |  | 29 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.826 | -9.522 | 1.849 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | min | [17 empty] | Sensor | REF | linear | 0.861 | 6.632 | 0.575 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.968 | -0.801 | 0.629 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.889 | -2.67 | 1.808 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.755 | 5.054 | 0.2685 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7645 | 4.901 | 0.2905 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7858 | 4.5891 | 0.3309 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7678 | 5.0321 | 0.2689 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7755 | 4.917 | 0.289 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7973 | 4.604 | 0.3303 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6279 | 6.69 | 0.221 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.666 | 6.384 | 0.244 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7091 | 6.023 | 0.282 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.653 | 3.891 | 0.345 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.621 | 4.692 | 0.355 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 4.2059 | 0.4089 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.659 | 4.3277 | 0.326 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6318 | 5.2236 | 0.33 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6659 | 4.8 | 0.3827 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6919 | 6.8046 | 0.601 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6933 | 6.7944 | 0.5973 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7224 | 5.9376 | 0.5485 |  |  |  |  | 98 |  | 21 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.798 | 8.2503 | 0.519 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7989 | 8.1951 | 0.5186 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.808 | 7.485 | 0.4769 |  |  |  |  | 98 |  | 21 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.249 | 23.604 | 0.9069 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2494 | 23.863 | 0.8989 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4076 | 16.185 | 0.888 |  |  |  |  | 98 |  | 31 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.1636 | 30.622 | 0.613 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.164 | 30.771 | 0.609 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.2519 | 24.712 | 0.6344 |  |  |  |  | 98 |  | 31 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/iqair-airvisual-pro---lab-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 10 | 5 | min | [17 empty] | Sensor | BAM | linear | 0.99 | 0.88 | 0.53 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -6.2 | 1140% | % | [21 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/iqair-airvisual-pro---lab-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.612 | 11.439 | 0.9303 |  |  |  |  | 100 |  | 8 | [38 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.55 | 41.991 | 1.064 |  |  |  |  | 100 |  | 3.8 | [38 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.609 | 12.972 | 0.7418 |  |  |  |  | 100 | [40 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.552 | 15.878 | 0.8323 |  |  |  |  | 100 | [40 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.52 | 10.394 | 0.461 |  |  |  |  | 77 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.563 | 11.136 | 0.619 |  |  |  |  | 77 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.54 | 11.04 | 0.392 |  |  |  |  | 77 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.533 | 12.385 | 0.5068 |  |  |  |  | 77 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6609 | 9.12 | 0.52 |  |  |  |  | 90 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.674 | 10.087 | 0.7 |  |  |  |  | 90 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.657 | 9.848 | 0.466 |  |  |  |  | 90 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 10.966 | 0.605 |  |  |  |  | 90 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.2307 | -0.1285 | 0.1301 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.6368 | -0.0562 | 0.057 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | counts | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7729 | -0.008 | 0.0493 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Met One |  |  | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | Met One-ES642 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | min |  |  | 0.75 | 0.25 |  |  |  | MLR | PM2.5=Rs*a1*(1+a2(RH2)/(1-RH)) |  |  | RH | [5 empty] | REF | Sensor | linear | 0.92 | [47 empty] | 5500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8489 | 9.12 | 0.731 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.884 | 9.274 | 0.8265 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.812 | 10.873 | 0.77 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8533 | 7.9017 | 0.793 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8707 | 8.2385 | 0.889 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8529 | 8.901 | 0.8785 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7 | 3.5 | 0.844 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.592 | 4.98 | 0.8647 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.744 | 4.228 | 0.985 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8 | 6.44 | 0.825 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.802 | 7.0186 | 0.919 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.781 | 7.369 | 0.909 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.759 | 3.31 | 0.37 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.658 | 4.72 | 0.858 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8 | 4.31 | 0.97 | [6 empty] | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS011 | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 2 | min | 0.8634 | 6.3 | 2 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | REF | Sensor |  |  |  |  | 2.6 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS012 | PM10 | µg/m3 | N | open source | LAB |  |  |  | 2 | min | 0.7679 | 5 | 0.7 |  |  |  |  | linear | PM10=a0+a1*Rs | [8 empty] | REF | Sensor |  |  |  |  | 3.3 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS013 | PM2.5 | µg/m3 | N | open source | FIELD | traffic |  |  | 24 | hour | [17 empty] | REF | Sensor | linear | 0.957 | 0.777 | 0.959 | 4.056 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS014 | PM10 | µg/m3 | N | open source | FIELD | traffic |  |  | 24 | hour | [17 empty] | REF | Sensor | linear | 0.909 | 1.288 | 0.972 | 6.084 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5711 | 5.883 | 0.4958 |  |  |  |  | 75 |  | 15 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5724 | 6.274 | 0.5414 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5862 | 6.2765 | 0.566 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0285 | 40.083 | 0.332 |  |  |  |  | 75 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0509 | 38.737 | 0.411 |  |  |  |  | 99.9 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.045 | 39.159 | 0.4191 |  |  |  |  | 99.9 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6636 | 6.3369 | 0.4788 |  |  |  |  | 75 |  | 15 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7664 | 6.0143 | 0.5643 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7498 | 6.2295 | 0.5746 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.01 | 43.548 | 0.1482 |  |  |  |  | 75 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0252 | 42.181 | 0.2293 |  |  |  |  | 99.9 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0258 | 42.083 | 0.2525 |  |  |  |  | 99.9 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1463 | 19.254 | 0.3471 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1479 | 1.7333 | 0.5978 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0026 | 32.344 | 0.0607 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.016 | 28.119 | 0.09 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.1714 | 0.6979 | 0.6428 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0023 | 32.472 | 0.0554 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0142 | 38.853 | 0.00958 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.1764 | 5.954 | 0.728 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0091 | 42.256 | 0.1243 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9384 | -0.9237 | 0.9738 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9369 | 0.3975 | 0.8723 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9467 | -0.0896 | 0.9071 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9025 | 2.3264 | 0.7228 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9087 | 3.0881 | 0.676 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9193 | 2.7581 | 0.6902 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3933 | 12.85 | 0.6651 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3969 | 13.378 | 0.6206 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.4124 | 12.932 | 0.6467 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9467 | -1.089 | 0.9825 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9405 | 0.3316 | 0.8757 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.952 | -0.1427 | 0.912 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9128 | 2.2696 | 0.7271 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9168 | 3.0625 | 0.6783 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9279 | 2.7253 | 0.6927 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4282 | 12.876 | 0.6646 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.432 | 13.416 | 0.6195 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4483 | 12.981 | 0.6448 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9884 | -0.8786 | 0.9695 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9834 | 0.4733 | 0.8623 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9884 | 0.054 | 0.894 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9676 | 2.4996 | 0.7134 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9682 | 3.3068 | 0.6635 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9739 | 3.0201 | 0.6747 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4779 | 13.357 | 0.6485 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4799 | 13.932 | 0.6017 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4962 | 13.542 | 0.6244 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7754 | 3.2227 | 0.7099 |  |  |  |  | 87 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7773 | 3.995 | 0.6637 |  |  |  |  | 87 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7944 | 3.6326 | 0.68 |  |  |  |  | 87 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9067 | 2.8604 | 0.7409 |  |  |  |  | 87 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9168 | 3.6259 | 0.6961 |  |  |  |  | 87 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9281 | 3.3161 | 0.7081 |  |  |  |  | 87 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3219 | 25.189 | 0.6845 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3264 | 25.704 | 0.6407 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3424 | 25.208 | 0.6708 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4578 | 24.94 | 0.7197 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4688 | 25.448 | 0.6773 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4864 | 25.01 | 0.7026 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.95 | -20.2 | 0.76 |  |  |  |  | 74.5 | 100 |  | 20 | 40 | [12 empty] | -4.1-77.2 over 15-300 ug/m3 | [23 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.997 | -2.41 | 0.35 |  |  |  |  | 100 | 100 |  | 20 | 40 | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.96 | -1.96 | 0.57 |  |  |  |  | 66 | 100 |  | 20 | 40 | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7593 | 3.5122 | 0.5884 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7493 | 4.4044 | 0.5785 |  |  |  |  | 99.5 |  | 11 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7585 | 4.1737 | 0.5652 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8621 | 3.4227 | 0.5845 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8479 | 4.3302 | 0.5717 |  |  |  |  | 99.5 |  | 11 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8595 | 4.1342 | 0.5562 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4563 | 9.4619 | 17.916 |  |  |  |  | 99.5 |  | 1.2 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.362 | 12.243 | 31.138 |  |  |  |  | 99.5 |  | 0.7 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4608 | 10.151 | 14.728 |  |  |  |  | 99.5 |  | 1.4 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9822 | 1.626 | 0.8125 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9849 | 2.0202 | 0.6981 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9784 | 1.4848 | 0.6969 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9655 | 2.3172 | 0.7873 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.855 | 1.8793 | 0.638 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9847 | 1.9355 | 0.8283 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9791 | 2.7288 | 0.6241 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9648 | 3.0028 | 0.5171 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9785 | 2.4318 | 0.5655 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9386 | 3.4081 | 0.5829 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9028 | 2.5956 | 0.5302 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9754 | 2.9308 | 0.636 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7037 | 7.3073 | 0.6882 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6785 | 7.7173 | 0.5325 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6967 | 6.9361 | 0.6099 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6677 | 8.0026 | 0.5985 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6598 | 7.0798 | 0.5871 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7014 | 7.4581 | 0.688 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9859 | 1.5968 | 0.8152 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9891 | 1.997 | 0.6996 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9818 | 1.4361 | 0.7 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9679 | 2.3088 | 0.7906 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9063 | 1.3309 | 0.6764 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.989 | 1.9103 | 0.8307 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9831 | 2.7136 | 0.6255 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9693 | 2.9952 | 0.5177 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9821 | 2.405 | 0.5678 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9416 | 3.4198 | 0.5849 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9384 | 2.2195 | 0.5511 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9801 | 2.9112 | 0.6376 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7212 | 7.3213 | 0.6907 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6966 | 7.7174 | 0.5346 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7065 | 7.0549 | 0.5987 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6851 | 8.0766 | 0.6008 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.699 | 6.6847 | 0.6082 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7235 | 7.4145 | 0.6922 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9928 | 1.5014 | 0.8266 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9918 | 2.0869 | 0.6964 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.99 | 1.1948 | 0.7185 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9745 | 2.529 | 0.7749 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9829 | 0.4563 | 0.7402 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9946 | 2.0312 | 0.8237 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9873 | 2.721 | 0.6305 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9716 | 3.1451 | 0.5152 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9873 | 2.2258 | 0.5808 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9489 | 3.7076 | 0.5756 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9859 | 1.7418 | 0.5812 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9854 | 3.0464 | 0.6345 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7731 | 6.9258 | 0.73 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7482 | 7.5023 | 0.559 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.003 | 0.4134 | -0.1364 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0013 | 0.4121 | -0.0723 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0025 | 0.4129 | -0.088 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0277 | 20.345 | -0.5731 |  |  |  |  | 100 |  | 0 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0207 | 14.535 | 0.1351 |  |  |  |  | 100 |  | 41 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0095 | 14.868 | 0.0674 |  |  |  |  | 100 |  | 79 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.402 | 15.173 | -0.1134 |  |  |  |  | 90 |  | 102 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3974 | 14.893 | 0.0002 |  |  |  |  | 96 |  | 100 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3381 | 14.002 | 0.0001 |  |  |  |  | 95 |  | 120 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4158 | 223.36 | 0.3731 |  |  |  |  | 93 |  | 157 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.8087 | 196.16 | 0.5864 |  |  |  |  | 100 |  | 157 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.7928 | 206.39 | 0.6051 |  |  |  |  | 90 |  | 117 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0439 | 7.1815 | 0.1555 |  |  |  |  | 93 |  | 10 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4568 | 2.673 | 1.2113 |  |  |  |  | 100 |  | 5 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.1103 | 3.6067 | 0.3914 |  |  |  |  | 90 |  | 10 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system |  | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0032 | 3.6916 | 0.0045 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system |  | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4393 | 1.9207 | 0.7121 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | updated | https://www.aqmesh.com/ | electrochemical | sensor system |  | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0075 | 3.4064 | 0.014 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.4855 | -0.0743 | -0.2945 | 1.3282 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.8303 | 0.171 | -0.8806 | 0.0481 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.7751 | -0.033 | 0.1717 | 0.871 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.487 | -0.0758 | -0.3303 | 1.3512 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.8437 | 0.1693 | 0.8764 | 0.0577 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7939 | 0.0034 | 0.6935 | 0.6935 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.4531 | -0.0571 | -0.5204 | 1.4293 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.8205 | 0.1187 | 0.5324 | 0.4258 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.771 | 0.0014 | 0.3249 | 0.7138 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0628 | 596.21 | -0.1267 |  |  |  |  | 5.5 |  | 1050 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8397 | 170.6 | 0.6855 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0751 | 779.71 | 0.2104 |  |  |  |  | 10.2 |  | 650 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8779 | 120.46 | 0.7236 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8726 | 170.33 | 0.6814 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8988 | 124.51 | 0.7126 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8926 | 132.59 | 0.688 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.9181 | 94.594 | 0.7484 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.1379 | 28.252 | -0.3172 |  |  |  |  | 96.7 |  | 24 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.1601 | 34.701 | -0.2835 |  |  |  |  | 99.5 |  | 55 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0004 | 6.1318 | -0.0094 |  |  |  |  | 0.4 |  | 17 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0586 | 27.758 | -0.1567 |  |  |  |  | 98.5 |  | 55 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/unitec-sens-it-co---lab-evaluation.pdf?sfvrsn=8 | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | LAB |  |  | 60 | 1 | min | [17 empty] | Sensor | FRM | linear | 0.9877 | -3.6185 | 3.0707 | [40 empty] | YES |  | YES |  |  | 2200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/unitec-sens-it-co---lab-evaluation.pdf?sfvrsn=8
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.426 | 0.1956 | -6e-05 | 0.343 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.15 | [14 empty] | 36.5 -50 over 0-25 ppm | [23 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.3431 | 0.2233 | 0.2085 | 0.1547 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.6 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.3299 | 0.2289 | 0.2031 | 0.1667 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.6 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.6017 | 2.2345 | 0.0012 | 0.1682 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 21 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.5678 | 2.2187 | 0.0012 | 0.1309 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 25 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.6207 | 1.2418 | 0.001 | 0.1689 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 26 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.628 | 2.2107 | 0.0011 | 0.1727 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 21 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.5977 | 2.2026 | 0.0012 | 0.1327 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 25 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.6475 | 1.217 | 0.0009 | 0.172 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 26 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.2832 | 0.2416 | 0.6047 |  |  |  |  | 71 |  | 0.1 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3136 | 0.2377 | 0.7756 |  |  |  |  | 71 |  | 0.07 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.2867 | 0.2433 | 0.572 |  |  |  |  | 71 |  | 0.12 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.00743 | 0.0104 | -0.0166 |  |  |  |  | 85 |  | 0.15 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.00717 | 0.0104 | -0.0205 |  |  |  |  | 85 |  | 0.14 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0821 | 0.0105 | -0.0142 |  |  |  |  | 85 |  | 0.16 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8026 | 0.233 | 0.9572 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8219 | 0.2417 | 0.8902 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8297 | 0.2316 | 1.0079 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.85 | 0.2315 | 0.9802 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8721 | 0.2377 | 0.9134 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8746 | 0.2305 | 1.0267 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8367 | 0.2432 | 0.9015 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8166 | 0.2491 | 0.8211 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.86123 | 0.2504 | 0.9105 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.5447 | -0.319 | 0.5114 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.4372 | 1.6073 | 0.3927 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.6051 | -1.7053 | 0.5 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.5643 | -0.3845 | 0.5203 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.4497 | 1.6573 | 0.3988 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.6225 | -1.8693 | 0.508 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.4275 | 4.3265 | 0.3246 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.2954 | 6.318 | 0.2212 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.4973 | 2.5653 | 0.345 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7669 | 6.3589 | 0.6559 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7338 | 8.0511 | 0.6212 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7833 | 5.7248 | 0.66886 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7796 | 7.1338 | 0.7249 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9194 | 2.3114 | 0.613 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8967 | 2.6344 | 0.5046 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9208 | 1.9955 | 0.5569 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | OEM | Envea - CairClip O3/NO2 | NO2 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.81 | 26.2 | 1.18 | -0.483 | -0.129 |  |  | MLR | Rs(NO2)=Rs-O3; NO2=a0+a1*Rs+a2*T+a3*RH |  | T | RH | [5 empty] | Sensor | FEM | linear | 0.57 | 11.391 | 0.955 | [6 empty] | 11.6 | [12 empty] | 40 | [20 empty] | YES |  | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AQMesh v.3.0 | CO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.75 | -0.0709 | 0.000888 | 0.002252 |  |  |  | MLR | CO=a0+a1*Rs+a2*day | [7 empty] | day | Sensor | FEM | linear | 0.63 | 0.0695 | 0.000809 | [6 empty] | 326.2 | [12 empty] | 0.7 | [25 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AQMesh v.3.0 | CO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.68 | 0.0512 | 0.000799 | [6 empty] | 326.2 | [12 empty] | 0.68 | [25 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO2-B42F (688150) | NO2 | ppb | N | black box | LAB |  | <25 |  | 15 | sec | 0.98 | -1.02 | 1.22 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | 10 | 63 | [8 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | CO-B4 (688150) | CO | ppb | N | black box | LAB |  | <15 |  | 15 | sec | 0.98 | 0.07 | 0.86 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.44 | 10 | 63 | [5 empty] | -0.034 |  |  | 25 | SD (ppb) | [25 empty] | 6 months - significant change | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO2-B42F (864150) | NO2 | ppb | N | black box | LAB |  | <25 |  | 15 | sec | 0.922 | 3.85 | 1.21 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | 10 | 63 | [8 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO2-B42F (688150) | NO2 | ppb | N | black box | FIELD | urban | <25 |  | 15 | sec | 0.423 | 3.8 | 0.38 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 30.27 | [5 empty] | 0.46 | [10 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | CO-B4 (688150) | CO | ppb | N | black box | FIELD | urban | <15 |  | 15 | sec | 0.336 | 166 | 0.88 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 170.99 | [5 empty] | 0.44 | [7 empty] | -0.034 |  |  | 25 | SD (ppb) | [25 empty] | 6 months - significant change | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO2-B42F (864150) | NO2 | ppb | N | black box | FIELD | urban | <25 |  | 15 | sec | 0.09 | 16 | 0.2 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | [10 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | OX-B421 (688150) | O3 | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | -1.27 | 1.16 |  |  |  |  | linear | [5 empty] | NO2 |  |  |  | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | OX-B421 (864150) | O3 | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | 3.25 | 0.99 |  |  |  |  | linear | [5 empty] | NO2 |  |  |  | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | OX-B421 (688150) | O3 | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.81 | 7.2 | 0.26 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 22.2 | [5 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | OX-B421 (864150) | O3 | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.32 | 9 | 0.11 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO-B4(688150) | NO | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | 0.07 | 0.86 |  |  |  |  | linear | [15 empty] | 16.35 | [44 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO-B4(688150) | NO | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.9801 | -1.13 | 0.97 |  |  |  |  | linear | [60 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
Envea |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | LAB |  | 2305 |  | 1 | hour | 0.9996 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.007 | -0.001 | -0.01 |  | 0.093 | -0.057 | 1.7 | SD | [13 empty] | 4.8 |  | 1.6 | < -0.065 | [9 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense | NO2 | ppb | Y | open source | LAB |  | 80 |  | 1 | hour | 0.9997 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.032 | < -1.3e-03 | 1.5 |  | 0.47 | 0.13 | 2.9 | SD | [13 empty] | 11.5 |  | 3.2 | < -0.497 |  |  |  |  | YES |  | YES |  |  | 180 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | Sensoric | NO2 | ppb | Y | open source | LAB |  | 98 |  | 1 | hour | 0.9992 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.0016 | -0.013 | 1.5 |  | 0.16 | 0.062 | 0.8 | SD | [13 empty] | 2.3 |  | 1.2 | < -0.196 | [9 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Envea |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | O3 | ppb | Y | open source | LAB |  | 150 |  | 1 | hour | 0.9973 | [12 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.84 | 0.006 | 0.07 |  |  | 0.24 |  | 3.3 | SD | [13 empty] | 9.2 |  | 1.4 | < 0.011 | [9 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | O3-B4 | Y | Y |  | electrochemical | OEM | AlphaSense | O3 | ppb | Y | open source | LAB |  | 84 |  | 1 | hour | 0.9998 | [12 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.926 | -0.042 | -0.066 |  |  |  | 0.4 | 0.4 | SD | [13 empty] | 0.9 |  | 0.7 | < 0.016 | [9 empty] | 274 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | sensoric | O3 | V vs ppb | Y | open source | LAB |  | 108 |  | 1 | hour | 0.9986 | 0.031 | 0.00485 | [10 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.76 | -0.011 | 7e-05 |  |  | 1.3 | -0.022 | 0.6 | SD | [13 empty] | 1.6 |  | 1.2 | < 0.142 | [9 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.108 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.009 | 522.71 | -73.28 | [45 empty] | 180 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.23 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.002 | 991.24 | -139.52 | [45 empty] | 180 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.351 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.026 | -108.36 | 17.07 | [45 empty] | 180 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.679 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.086 | -33.59 | 6.96 | [45 empty] | 180 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | sensoric | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.001 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.068 | 1727 | -247.53 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | sensoric | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.002 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.51 | -1977.94 | 280.68 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | sensoric | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.768 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.078 | -78.28 | 1.097 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | sensoric | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.562 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.062 | -108.18 | 16.01 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.206 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.131 | 8.13 | -58.83 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.2 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.126 | 8.48 | -56.45 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.168 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.016 | 56.37 | -378.51 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.269 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.203 | 6.18 | -40 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.745 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a0 | O3 | T | [6 empty] | REF | Sensor | linear | 0.021 | 3176.44 | -477.45 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.744 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a0 | O3 | T | [6 empty] | REF | Sensor | linear | 0.063 | -207890.74 | 31407.64 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.525 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*NO+a4*T+a0 | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | 0.01 | -3970.53 | 577.84 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.786 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*NO+a4*T+a0 | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | 0.016 | 5692.9 | -852.21 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.465 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.004 | 438.99 | -60.94 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.24 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.036 | 229.86 | -29.14 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.745 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.021 | 16.17 | -97.84 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.585 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.004 | 32.5 | -205.92 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-B4 | Y | N |  | electrochemical | OEM | O3-B4 | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.023 | [6 empty] | linear | Rs=a1*O3+a0 | [8 empty] | REF | Sensor | linear | 0.212 | -142.99 | 4.83 | [45 empty] | 274 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-B4 | Y | N |  | electrochemical | OEM | O3-B4 | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.503 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a3*NO2+a4*H2O+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.479 | -37.97 | 2.33 | [45 empty] | 274 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | sensoric | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.845 | [6 empty] | linear | Rs=a1*O3+a0 | [8 empty] | REF | Sensor | linear | 0.667 | -26.31 | 1.84 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | sensoric | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.8478 | [6 empty] | linear | Rs=a1*O3+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.813 | -20.29 | 1.65 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | sensoric | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.852 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.584 | -30.15 | 2.1 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | sensoric | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.945 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.824 | -12.03 | 1.49 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | electrochemical | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.3 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.136 | 2.24 | 13.1 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | electrochemical | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.17 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.151 | 3.39 | 17.7 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | electrochemical | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.27 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.022 | -11.08 | 57.6 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | electrochemical | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.3 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.036 | 8.15 | -39 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.78 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.066 | -1.48 | 9.64 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.78 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.067 | -1.76 | 11.03 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.76 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.035 | -2.92 | 17.19 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2015 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.77 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.047 | -2.59 | 15.19 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | sensoric | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.0057 | [6 empty] | linear | Rs = a1*NO+a0 | [8 empty] | REF | Sensor | linear | 0.018 | -315.8 | 307 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | sensoric | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 2.3e-05 | [6 empty] | linear | Rs = a1*NO+a0 | [8 empty] | REF | Sensor | linear | 0.001 | -21860 | 23106 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | sensoric | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.0018 | [6 empty] | MLR | Rs = a1*NO+a2*T+a3*RH+a0 | [8 empty] | REF | Sensor | linear | 0.009 | 972 | -1056.6 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | sensoric | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.075 | [6 empty] | MLR | Rs = a1*NO+a2*T+a3*RH+a0 | [8 empty] | REF | Sensor | linear | 0.02 | 118.62 | 108.9 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Perkin Elmer |  |  | US-EPA_2015 | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | ELM | N | N | https://scistarter.com/tools/776 | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  |  | 1 | hour | [17 empty] | CAPS | Sensor | linear | 0.004 | 18.552 | 0.808 | [8 empty] | > 92% (spike) | [36 empty] | 5200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8682 | 3.0635 | 0.5693 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8791 | 1.8201 | 0.5404 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9162 | 2.5066 | 0.6248 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 11.207 | 0.7476 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6247 | 11.701 | 0.5753 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6505 | 10.81 | 0.6641 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6067 | 12.161 | 0.6427 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6379 | 10.485 | 0.6598 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6536 | 11.332 | 0.7479 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9713 | 2.4968 | 0.6057 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9528 | 2.919 | 0.4941 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9721 | 2.0167 | 0.5582 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9269 | 3.4751 | 0.551 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9657 | 1.5789 | 0.5571 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6092 | 2.8139 | 0.6092 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7336 | 11.161 | 0.7717 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7017 | 11.844 | 0.5873 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7313 | 10.527 | 0.6949 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6827 | 12.469 | 0.65 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7475 | 9.85 | 0.7086 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7359 | 11.415 | 0.7641 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 1 | -8.49 | 1.49 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 85-95 over 0-175 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -15.64 | 1.31 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 54-96 over 0-250 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.94 | -19.79 | 2.38 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 44-73 over 21-180 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9972 | [16 empty] | 19 ppb NO2 at > 200 ppb SO2 |  |  | 1.2 | SD | [12 empty] | ppb | [13 empty] | 0 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9919 | [19 empty] | 6.4 | SD | [12 empty] | ppb | [13 empty] | 0 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9937 | [19 empty] | 7.4 | SD | [12 empty] | ppb | [13 empty] | 0 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9917 | [19 empty] | 7.5 | SD | [12 empty] | ppb | [13 empty] | 0 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9968 | [28 empty] | 7 |  |  |  | 3.8 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9978 | [28 empty] | 0 |  |  |  | 2.9 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9975 | [28 empty] | 8.6 |  |  |  | 4.2 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9896 | [28 empty] | 8.3 |  |  |  | 9.3 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Alphasense |  | CitiSense | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NO2-A1 | Y | Y |  | electrochemical | OEM |  | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.9772 | [16 empty] | 34.2ppb @ > 200ppb SO2 |  |  | 4.6 | SD | [21 empty] | YES |  |  |  |  | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NORMAL | JRC-FK
Alphasense |  | CitiSense | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NO2-A1 | Y | Y |  | electrochemical | OEM |  | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.9722 | [16 empty] | 34.2ppb @ > 200ppb SO2 |  |  | 23.3 | SD | [21 empty] | YES |  |  |  |  | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HUMID | JRC-FK
Platypus |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | Platypus NO2 | N | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.7956 | [32 empty] | 18 | ppb | [13 empty] | 50 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
SGX-Sensortech |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | MiCS-2710 | Y | N |  | MOs | OEM | SGX Sensortech - MiCS-2710 | NO2 | ppb | N | open source | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9851 | [19 empty] | 5 | SD | [12 empty] | ppb | [8 empty] | YES |  |  |  |  | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HOT | JRC-FK
SGX-Sensortech |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | MiCS-2710 | Y | N |  | MOs | OEM | SGX Sensortech - MiCS-2710 | NO2 | ppb | N | open source | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9815 | [19 empty] | 2.6 | SD | [12 empty] | ppb | [8 empty] | YES |  |  |  |  | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HUMID | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8085 | 5.5525 | 0.487 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | LAB |  |  |  | 1 | hour | 0.92 | 115 | 0.72 |  |  |  |  | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.92 | 26 | 0.89 |  |  |  |  | MLR | Rs=a1*CO+a2*T+a3*RH |  | T | RH | [5 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.91 | 46 | 0.89 |  |  |  |  | RF | [9 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | LAB |  |  |  | 1 | hour | 0.31 | 34.8 | 0.3 |  |  |  |  | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | [21 empty] | 3.5 ppb | [28 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.27 | 8.8 | 0.33 |  |  |  |  | MLR | Rs=a1*NO2+a2*T+a3*RH |  | T | RH | [5 empty] | REF | Sensor | [50 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.62 | 4.5 | 0.72 |  |  |  |  | RF | [9 empty] | REF | Sensor | [50 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8003 | 8.6212 | 0.5092 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8684 | 4.3356 | 0.5592 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8205 | 5.4635 | 0.4924 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8102 | 8.2603 | 0.5227 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8762 | 4.3305 | 0.5548 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8436 | 5.3419 | 0.5 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7289 | 6.7195 | 0.8546 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8967 | 4.0953 | 0.566 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com | Intel Berkeley Badge | Y | N |  | electrochemical | sensor system | City Technology - NO2 3E50 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.9862 | 1622.4 | 0.0857 | [14 empty] | REF | Sensor | linear | [49 empty] | EUR | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com | Intel Berkeley Badge | Y | N |  | electrochemical | sensor system | City Technology - NO2 3E50 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.9823 | 1628 | 0.0948 | [14 empty] | REF | Sensor | linear | [49 empty] | EUR | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com | Intel Berkeley Badge | Y | N |  | MOs | sensor system | SGX Sensortech - MICS-4514 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.8913 | 690.16 | 488.47 | [14 empty] | REF | Sensor | logaritmic | [49 empty] | USD | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com |  | JRC-FK
Aeroqual |  |  | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com | S-500 | Y | Y |  | MOs | sensor system |  | NO2 | ppb | Y | black box | LAB |  | < 60 |  | 1 | hour | 0.8008 | -2.2563 | 0.8899 | [14 empty] | REF | Sensor | linear | [25 empty] | 15% |  | 1 | ppb |  |  | 0-0.2 | ppm | [15 empty] | 500 | USD | Vaughn, D. L., Dye, T. S., Roberts, P. T., Ray, A. E., & DeWinter, J. L. (2010). Characterization of low-Cost NO2 Sensors (No. Characterization of Low-Cost NO 2 Sensors). U.S. Environmental Protection Agency Research Triangle Park, NC. Retrieved from sonomatech.com |  | JRC-FK
Alphasense |  |  | Sun et al. 2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1916 | NO2-B43F | Y | Y |  | electrochemical | OEM |  | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | [7 empty] | MLR | NO2 = a1*Rs +a2*T +a3*RH +a4 |  | T | RH | [5 empty] | Sensor | REF | linear | 0.99 | 0 | 0.95 | [45 empty] | 280 | USD | https://www.mdpi.com/1424-8220/17/8/1916 |  | JRC-FK
Aeroqual |  |  | Lin et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231014008607 | S-500 | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [7 empty] | linear | Rs=NO2+a1*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.88 | -1.1 | 1.03 | [45 empty] | 500 | USD | https://www.sciencedirect.com/science/article/pii/S1352231014008607 |  | JRC-FK
TSI |  |  | Karagulian et al. 2012 | 2012 | https://pubs.rsc.org/en/content/articlelanding/2012/em/c2em30099k#!divAbstract | SidePak AM510 | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | rural |  |  | 30 | min | [17 empty] | TEOM-FDMS | Sensor | linear | 0.94 | 0 | 1 | [36 empty] | YES |  | YES |  | YES |  |  |  |  | 3000 | USD | https://pubs.rsc.org/en/content/articlelanding/2012/em/c2em30099k#!divAbstract |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7809 | 6.8335 | 0.5547 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6743 | 9.7879 | 0.5939 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.801 | 5.7815 | 0.6133 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8956 | 6.4465 | 0.5919 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7733 | 9.2457 | 0.6421 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9105 | 5.1977 | 0.6545 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---lab-evaluation.pdf | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | 9.41 | 0.31 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -27- (-101.5) over 0-240 ug/m3 | [23 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---lab-evaluation.pdf |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.573 | 3.778 | 1.1198 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7951 | -0.0068 | 1.2038 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7186 | -8.491 | 1.1094 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8713 | -7.5874 | 1.496 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8758 | -2.1549 | 1.3772 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8036 | -1.4041 | 1.1638 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7294 | -9.3819 | 1.4734 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | hour | 0.4 | 3.7 | 0.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 2.01 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | hour | 0.41 | 4.5 | 0.7 | -0.06 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 2.07 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 6 | hour | 0.8 | -1.9 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.53 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 12 | hour | 0.84 | -2.4 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.35 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 24 | hour | 0.93 | -4.2 | 1.6 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.15 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | min | 0.66 | -2.1 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 27 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | min | 0.93 | -3.5 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | min | 0.94 | -1.5 | 1.5 | -0.09 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | hour | 0.66 | -2.1 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 27 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | hour | 0.93 | -3.5 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 10 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 1 | hour | 0.95 | -1.4 | 1.5 | -0.09 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 8 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 6 | hour | 0.73 | -2.4 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 22 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 6 | hour | 0.92 | -3.2 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 10 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 6 | hour | 0.95 | -0.3 | 1.5 | -0.13 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 8 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 12 | hour | 0.84 | -3.4 | 1.9 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 15 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 12 | hour | 0.93 | 8.7 | 2.2 | -0.49 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | suburban |  |  | 24 | hour | 0.94 | -4.1 | 2 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.61 | 5.1 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 46 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.75 | 74 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 35 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.74 | 74 | 1.4 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 24 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.61 | 88 | 1.2 | -2.3 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 46 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.78 | 276 | 1.1 | -7.9 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 30 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 6 | hour | 0.8 | -5.8 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 32 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 6 | hour | 0.87 | 65 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 25 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 6 | hour | 0.86 | 65 | 1.4 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 16 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 6 | hour | 0.81 | 47 | 1.6 | -1.4 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 30 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 6 | hour | 0.9 | 248 | 1.2 | -7 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 18 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 12 | hour | 0.84 | -6.5 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 18 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 12 | hour | 0.88 | 74 | 1.3 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 19 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 12 | hour | 0.86 | 74 | 1.3 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 12 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 12 | hour | 0.89 | 330 | 1 | -10 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 17 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 24 | hour | 1 | -4.5 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 17 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 24 | hour | 0.99 | 82 | 1.2 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 24 | hour | 1 | 82 | 1.2 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.28 | 1.91 | 1.16 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.01 | 13.6 | 0.04 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.13 | 10.92 | 1.28 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | NO2 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.28 | 24.23 | 1.57 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.79 | 4.79 | 0.93 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.72 | 3.35 | 1.1 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.471 | 455.38 | 0.6977 | -1.0835 |  |  |  | MLR | NO2=a0+a1*Rs1+a2*Rs2 | [8 empty] | REF | Sensor | linear | 0.288 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.612 | 715.45 | 0.8394 | -1.4811 |  | 0.5326 |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a4*RH |  |  | RH | [5 empty] | REF | Sensor | linear | 0.293 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.556 | 827.92 | 0.8688 | -1.5498 | -1.6344 |  |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T |  | T | [6 empty] | REF | Sensor | linear | 0.821 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.617 | 790.88 | 0.8707 | -1.5645 | -0.5051 | 0.4513 |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T+a4*RH |  | T | RH | [5 empty] | REF | Sensor | linear | 0.82 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.788 | 274.85 | 0.3186 | -0.4805 | -0.5447 | -0.4744 | -0.5349 | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T+a4*RH+a5*O3 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.887 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban | 60 |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  | 0.81 | [11 empty] | 0.012 | -0.0002 | 1 | [19 empty] | 0.94 | [12 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban | 60 |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 |  | 0.81 | [11 empty] | 0.012 | -0.002 | 1 | [19 empty] | 0.94 | [12 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
Shinyei |  |  | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | PPD42NS | Y | Y |  | nephelometer | OEM | PPD42NS | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2.5 | hour | 0.9452 | -0.353 | 33.6 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 20 | [5 empty] | 6.44 | 4.59 | [8 empty] | YES | 300 | YES | 600 | YES | 400 | YES | 0.15 |  | 15 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.6246 | -2.1884 | 1.619 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.572 | -2.8132 | 1.2934 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.94 | 0.8 | 0.94 |  |  |  |  | linear | Rs=a0+a1*NO2 | [8 empty] | REF | Sensor | linear | 0.69 | 3.26 | 0.81 | 4.56 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.96 | 9.19 | 0.96 |  |  |  |  | linear | Rs=a0+a1*CO | [5 empty] | CO |  |  | REF | Sensor | linear | 0.88 | 3.98 | 0.94 | 32.9 | [9 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - NO-B4 | NO | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.94 | 0.46 | 0.94 |  |  |  |  | linear | Rs=a0+a1*NO | [8 empty] | REF | Sensor | linear | 0.84 | 1.29 | 0.94 | 4.52 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - OX-B421 | O3 | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.65 | 8.34 | 0.65 |  |  |  |  | linear | Rs=a0+a1*O3 | [8 empty] | REF | Sensor | linear | 0.39 | 13.1 | 0.47 | 9.71 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | nephelometer | sensor system | PPD20V | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.33 |  |  |  | 21.15 | [43 empty] | 30 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | OPC | sensor system | CAIR | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.13 |  |  |  | 80.91 |  |  |  | 93 | [39 empty] | 200 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7133 | 4.2246 | 0.4975 |  |  |  |  | 80 | 100 | 14 | [14 empty] | 35- 69 over 16-270 ug/m3 | [23 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | Shinyei -shinyei-PPD42 | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.33 |  |  |  | 24.57 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | Shinyei -shinyei-PPD42 | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.36 |  |  |  | 23.61 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | OPC | sensor system | CAIR | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.07 |  |  |  | 17.15 | [43 empty] | 200 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | shinyei-PPD42 | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.23 |  |  |  | 11.26 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | shinyei-PPD42 | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.27 |  |  |  | 11.11 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.84 |  |  |  | 5.61 |  |  |  | 93 | [39 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.06 |  |  |  | 20.17 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Ingenieros Assessores | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.57 |  |  |  | 14.93 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | electrochemical | sensor system | CityTechnology-NO2 3E50 | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  |  |  | 4.95 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  |  |  | 1.46 | [43 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
SGX-Sensotech |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ISAG | N | N |  | MOs | sensor system | MiCS-2710 | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.02 |  |  |  | 16.17 | [43 empty] | 300 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | electrochemical | sensor system | MiCS-2710 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.76 |  |  |  | 0.09 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Ingenieros Assessores | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | N | Y |  | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.53 |  |  |  | 0.1 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 170 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 |  |  |  | 0.18 |  |  |  | 92 |  |  |  |  | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 170 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.86 |  |  |  | 0.05 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | electrochemical | sensor system | Alphasense - O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.13 |  |  |  | 22.12 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Ingenieros Assessores | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | Y | Y |  | MOs | sensor system | SGX-Sensotech-MICS-OZ-47 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.77 |  |  |  | 7.66 | [43 empty] | 0 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | electrochemical | sensor system | Alphasense-O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.14 |  |  |  | 21.5 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7 |  |  |  | 2.4 | [43 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Aristotle University |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ISAG | N | N |  | MOs | sensor system | SGX-Sensotech - MICS-2610 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.12 |  |  |  | 360.12 | [43 empty] | 300 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Alphasense |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | AlphaSense NO-A4 | NO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [20 empty] | 0.34 |  |  |  | 12.02 |  |  |  | 93 | [39 empty] | 165 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system | AlphaSense NO-A4 | NO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [20 empty] | 0.8 |  |  |  | 1.51 |  |  |  | 37 | [39 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7277 | 3.9572 | 0.4781 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7425 | 2.8438 | 0.566 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7761 | 4.4018 | 0.4855 |  |  |  |  | 80 | 100 | 14 | [14 empty] | 35- 69 over 16-270 ug/m3 | [23 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7943 | 4.0539 | 0.4727 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
AQmesh |  | POD1 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.81 | 11.542 | 0.727 | 0.097 | -0.437 | 1.5416 |  | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | MLR (R), Support Vector Machines, e1017 (R), ANN - Neuralnet (R), RandomForest (R), R caret PACKAGES | JRC-FK
AQmesh |  | POD2 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.93 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD3 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.85 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD4 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.9 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD1 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.9 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.91 |  |  | 4.88 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD2 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.86 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 8.67 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD3 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.85 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.91 |  |  | 4.58 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD4 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.95 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.97 |  |  | 7.44 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD1 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.81 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.92 |  |  | 4.09 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD2 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.89 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.95 |  |  | 8.26 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD3 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.79 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.92 |  |  | 4.16 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD4 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.95 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.97 |  |  | 5.68 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD1 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.88 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.36 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD2 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.62 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.98 |  |  | 6.34 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD3 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.88 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.4 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  | POD4 | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | updated |  | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.86 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.93 |  |  | 7.09 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -15.91 | 0.1593 |  | 0.7 |  |  | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.9606 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.84 | 10 | 0.82 | 7.96 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -15.93 | 0.1606 | -0.00175 | 0.7 | 0.000121 | 1.87e-07 | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8836 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.85 | 12.6 | 0.84 | 10.33 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -0.09 | 0.1547 |  | 0.63 |  |  | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8464 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.87 | 7.4 | 0.85 | 6.14 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -0.13 | 0.1548 | -0.000204 | 0.63 | 0.000134 | 1.81e-07 | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8464 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.8 | 11.1 | 0.83 | 9.23 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8276 | 2.8625 | 0.5641 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/sainsmart---lab-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 30 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -4.13 | 0.79 |  |  |  |  | 99.6 | 100 |  | 20 | 40 | [36 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/sainsmart---lab-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3116 | -1.1229 | 0.5626 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0006 | 12.106 | -0.0029 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.2708 | -0.5655 | 0.497 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1206 | 28085 | 211.16 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Envea |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7569 | -10 | 0.65 |  |  |  |  | 63 | [40 empty] | 600 | EUR | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Envea |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | OEM | CairPol | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7056 | -15 | 0.67 |  |  |  |  | 82 | [40 empty] | 600 | EUR | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0008 | 147465 | 2.0367 |  |  |  |  | 95 |  | 980 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0532 | 52685 | 144.9 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3288 | -2.0194 | 0.6009 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0008 | 12.127 | -0.0036 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2844 | -1.3926 | 0.5326 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.1264 | 19465 | 226.47 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0009 | 147152 | 2.2379 |  |  |  |  | 95 |  | 980 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.053 | 48225 | 152.68 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2486 | -0.8308 | 0.6312 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 7e-06 | 13.908 | -0.0004 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.1482 | 1.678 | 0.472 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Samyoung |  | Samyoung | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | DSM501A | Y | Y |  | nephelometer | OEM | DSM501A | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2.5 | hour | 0.8914 | 3.93 | 59.7 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 19 | [5 empty] | 11.4 | 4.28 | [8 empty] | YES | 100 | YES | 250 | YES | 100 | YES | 15 |  | 15 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
Samyoung |  | Samyoung | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | DSM501A | Y | Y |  | nephelometer | OEM | DSM501A | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 24 | hour | 0.5 | [6 empty] | linear | [9 empty] | Sensor | DustTrak | [49 empty] | 15 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Sensaris |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Eco PM | N | N |  | OPC | sensor system |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 30 | sec | [17 empty] | GRIMM | Sensor | linear | 0.315 | 0.0037 | 0.034 | [45 empty] | 0 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Sharp |  |  | Olivares and Edwards, 2015 | 2015 | https://www.atmos-meas-tech-discuss.net/amt-2015-175/ | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [7 empty] | linear | PM2.5=0.6*Rs+19; Rs=-4.4*T - 49 |  | T | [6 empty] | TEOM-FDMS | Sensor | linear | 0.72 |  |  | 14.1 | [41 empty] | NO |  |  | 10 | USD | https://www.atmos-meas-tech-discuss.net/amt-2015-175/ |  | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.98 | 3.24 | 0.45 |  |  |  | 46.1 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.94 | 3.4 | 0.3 |  |  |  | 96.9 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.98 | 1.56 | 0.41 |  |  |  | 118.2 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -55 | 1.3 |  |  | -5.8 | [16 empty] | 3.1 | [16 empty] | YES | 300 | [7 empty] | 10 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Sharp |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | DN7C3CA006 | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -56 | 1.3 |  |  | -6.3 | [16 empty] | 4.2 | [16 empty] | YES | 120 | [7 empty] | 10 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Sharp |  |  | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | mg/m3/Volts | Y | open source | LAB |  |  |  | 1 | sec | 0.9844 |  | -4e-05 | 0.2227 |  |  |  | quad | Rs=a1*PM2.52+a2Rs | [8 empty] | Sensor | DustTrak | [49 empty] | 10 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Sharp |  |  | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM10 | mg/m3/Volts | Y | open source | LAB |  |  |  | 1 | sec | 0.6461 |  | -0.0002 | 0.2908 |  |  |  | quad | Rs=a1*PM102+a2Rs | [8 empty] | Sensor | DustTrak | [49 empty] | 10 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8267 | 4.123 | 0.6655 |  |  |  |  | 99 |  | 16 | [14 empty] | -70 - 91.2 over 0-180 ug/m3 | [23 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8152 | 4.1645 | 0.6718 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8274 | 4.202 | 0.7498 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9299 | 3.5897 | 0.6942 |  |  |  |  | 99 |  | 16 | [14 empty] | -70 - 91.2 over 0-180 ug/m3 | [23 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9245 | 3.4702 | 0.7101 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9323 | 3.6368 | 0.7838 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8971 | 3.0295 | 0.578 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8896 | 2.988 | 0.5956 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.889 | 3.1992 | 0.6448 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9183 | 2.8934 | 0.5874 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9119 | 2.836 | 0.6061 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9148 | 3.0078 | 0.6587 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9509 | 2.6128 | 0.603 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6284 | 2.426 | 0.6284 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6792 | 2.6785 | 0.6792 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---lab-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.93 | -18.46 | 0.84 | [7 empty] | 20 | 40 | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---lab-evaluation.pdf | PM evaluation Kit | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system | Shinyei PMS-SYS-1 | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.42 | 11.3 | 0.78 | -0.04 | -0.0112 |  |  | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day |  |  | RH |  |  |  |  | day | Sensor | FEM | linear | 0.36 | 0.748 | 0.72 | [6 empty] | 12.18 | [12 empty] | 0.69 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-1 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AQMesh v.3.0 | NO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.75 | [6 empty] | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day | [8 empty] | Sensor | FEM | linear | 0.77 | 1.06 | 0.892 | [6 empty] | 18.3 | [12 empty] | 1.84 | [22 empty] | YES |  |  | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE-2 |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AQMesh v.3.0 | NO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.87 | -0.765 | 0.883 | [6 empty] | 18.5 | [12 empty] | 1.76 | [22 empty] | YES |  |  | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
Shinyei |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0033 | 5.2302 | 0.0659 | [8 empty] | > 95% (spike) | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Shinyei |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.1516 | 2.756 | 0.2917 | [8 empty] | > 95% (spike) | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5041 | 0.24 | 0.58 | 0.2 |  |  |  | 73 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5184 | 0.8 | 0.54 |  |  |  |  | 87 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0001 | 4.35 | 0.42 |  |  |  |  | 92 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.92 | -35 | 1.6 |  |  |  | 22.3 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.92 | -35 | 1.8 |  |  |  | 22.3 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.96 | 2.9 | 2 |  |  |  | 52.7 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | hour | 0.58 | [6 empty] | linear | PM2.5=a0+a1*Rs | [8 empty] | BAM | Sensor | [44 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | hour | 0.59 | [6 empty] | MLR | PM2.5=a0+a1*Rs+a2*RH |  |  | RH | [5 empty] | BAM | Sensor | [44 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.92 |  |  | 2.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.94 |  |  | 2 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.94 |  |  | 2.1 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.9 |  |  | 1.7 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.93 |  |  | 1.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.93 |  |  | 1.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.64 |  |  | 5.1 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.68 |  |  | 4.8 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.7 |  |  | 4.6 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.56 |  |  | 3.5 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.61 |  |  | 3.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.58 |  |  | 3.6 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Austin et al. 2015 | 2015 | http://dx.plos.org/10.1371/journal.pone.0137789 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2 | min | [7 empty] | exp | Rs=a0(1-exp(-K*PM2.5)) | [8 empty] | APS | Sensor | linear | 0.97 |  | 5.4 | [27 empty] | 1 | [10 empty] | YES | [6 empty] | 15 | USD | http://dx.plos.org/10.1371/journal.pone.0137789 | Arizona Test Dust | JRC-FK
Shinyei |  |  | Gao et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115000160 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | urban |  |  | 1 | min | [7 empty] | quad |  |  | T | RH | [5 empty] | DustTrak | Sensor | linear | 0.925 | [42 empty] | YES |  | YES |  |  | 15 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115000160 |  | JRC-FK
Shinyei |  |  | Gao et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115000160 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | urban |  |  | 1 | hour | [7 empty] | quad |  |  | T | RH | [5 empty] | DustTrak | Sensor | linear | 0.885 | [42 empty] | YES |  | YES |  |  | 15 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115000160 |  | JRC-FK
Shinyei |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 10 | min | [17 empty] | REF | Sensor | linear | 0.65 | [47 empty] | 15 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.96 | 5.3 | 2.3 |  |  |  | 52.7 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.29 | 6.7 | 0.74 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.21 | 5.4 | 1.06 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.07 | 5.9 | 0.4 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.28 | 7.7 | 0.55 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.46 | 4.3 | 1.9 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.15 | 5.7 | 0.66 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.82 | -1370 | 1.35 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.89 | -3799 | 1.81 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.72 | -1300 | 1.3 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.85 | -2299 | 1.61 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.75 | -432 | 1.16 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8137 | 2.589 | 0.6731 |  |  |  |  | 99 |  | 16 | [14 empty] | -108- (-7.2) over 0-150 ug/m3 | [23 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8304 | 0.9655 | 0.6959 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.81 | 0.9805 | 0.7248 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8299 | 2.4787 | 0.6797 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8466 | 0.8409 | 0.7033 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8297 | 0.7892 | 0.736 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8716 | 2.629 | 0.6768 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8779 | 0.9703 | 0.7064 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8632 | 0.9374 | 0.7403 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7647 | 2.1327 | 0.6852 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7734 | 0.5203 | 0.7062 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.757 | 0.4767 | 0.7394 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8737 | 2.1322 | 0.6705 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 10 | hour | [17 empty] | REF | Sensor | linear | 0.84 | [29 empty] | 3.22 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.845 | 5.22 | 1.26 |  |  |  |  | linear | Rs=a0+a1*PM2.5 | [8 empty] | TEOM | Sensor | [31 empty] | 10.5 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.92 | -7.16 | -0.033 | 90.9 |  |  |  | exp | Rs= a0+a1*(1-exp(a2*PM2.5)) | [8 empty] | TEOM | Sensor | [31 empty] | 10.5 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS3003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 10 | hour | [17 empty] | REF | Sensor | linear | 0.85 | [47 empty] | 30 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
City University of Hong Kong |  | HKEPD/14-02771 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppm/mV | Y | black box | LAB |  |  |  | 1 | min | 0.99 | -15.14 | 1.08 |  |  |  |  | linear | CO=a0+a1*Rs | [8 empty] | Sensor | REF | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [25 empty] | NO |  | NO |  |  |  | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
City University of Hong Kong |  | HKEPD/14-02772 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - NO2-B4 | NO2 | ppb/mV | Y | black box | LAB |  |  |  | 1 | min | 0.99 | -0.99 | 0.0026 |  |  |  |  | linear | NO2=a0+a1*Rs | [8 empty] | Sensor | REF | [44 empty] | YES | [5 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
City University of Hong Kong |  | HKEPD/14-02773 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppm/mV | Y | black box | FIELD | urban |  |  | 1 | min |  | 0.099 | 0.0025 |  |  |  |  | linear | CO=a0+a1*Rs | [8 empty] | REF | Sensor | linear | 0.97 | [12 empty] | 0.085 |  |  | -0.053 | -0.034 | [25 empty] | NO |  | NO |  |  |  | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
City University of Hong Kong |  | HKEPD/14-02774 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - NO2-B4 | NO2 | ppb/mV | Y | black box | FIELD | urban |  |  | 1 | min |  | -3.96 | 10.97 | -0.17 | 0.33 |  |  | MLR | NO2=(Rs-a1*RH-a0)/(a2*RH+a3) |  |  | RH | [5 empty] | REF | Sensor | linear | 0.9 | [42 empty] | YES | [5 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
2B Technologies |  | HKEPD/14-02774 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N | https://twobtech.com/products.html | UV | sensor system | 2B Tech. (POM) | O3 | ppb/mV | Y | open source | LAB |  |  | 20 | 1 | min | 0.99 | -0.99 | 0.0026 |  |  |  |  | linear | O3=a0+a1*Rs | [8 empty] | Sensor | REF | [50 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8853 | 0.3528 | 0.7035 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.99 | [12 empty] | 0.2 | 0.24 | [34 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | good sensor-to-sensor reproducibility | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - NO-AF | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.8 | [12 empty] | 0.0045 |  | 0.001 | [33 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - NO-AF | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.95 | [12 empty] | 0.0045 |  | 0.001 | [33 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | LAB |  |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a1*exp(a2*CO)*(T-298)+a3*exp(a4*CO)+a5*exp(a6*CO)*RH) |  | T | RH | [11 empty] | 3.56 | [32 empty] | 0.49 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a0+a1*CO+a2*T+a3*RH+a4*time) |  | T | RH | [11 empty] | 0.44 | [32 empty] | 0.49 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | LAB |  |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a1*exp(a2*NO2)*(T-298)+a3*exp(a4*NO2)+a5*exp(a6*NO2)*RH) |  | T | RH | [11 empty] | 7.1 | [32 empty] | 0.93 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a0+a1*NO2+a2*T+a3*RH+a4*time) |  | T | RH | [11 empty] | 8.4 | [32 empty] | 0.93 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.16 | [8 empty] | 307.1 | [12 empty] | 0.31 | [26 empty] | USD | https://www.atmos-meas-tech.net/7/3325/2014/ |  | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | Y | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.06 | [8 empty] | 11.3 | [12 empty] | 2.46 | [26 empty] | USD | https://www.atmos-meas-tech.net/7/3325/2014/ |  | JRC-FK
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8876 | 0.2103 | 0.7443 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | N |  | electrochemical | OEM |  | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.96 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.09 | SD | [21 empty] | YES |  | YES |  |  | 170 | USD | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | N |  | electrochemical | OEM |  | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.98 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.03 | SD | [21 empty] | YES |  | YES |  |  | 170 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | N |  | electrochemical | OEM |  | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.97 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.03 | SD | [21 empty] | YES |  | YES |  |  | 170 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | N |  | electrochemical | OEM |  | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.98 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.02 | SD | [21 empty] | YES |  | YES |  |  | 170 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM |  | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.84 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 6.5 | SD | [21 empty] | YES |  | YES |  |  | 280 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM |  | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.79 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 5.8 | SD | [21 empty] | YES |  | YES |  |  | 280 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM |  | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 7.5 | SD | [21 empty] | YES |  | YES |  |  | 280 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM |  | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 5.8 | SD | [21 empty] | YES |  | YES |  |  | 280 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM |  | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.7 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 4.5 | SD | [21 empty] | YES |  | YES |  |  | 274 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM |  | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.73 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 6.1 | SD | [21 empty] | YES |  | YES |  |  | 274 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM |  | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.73 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 13.4 | SD | [21 empty] | YES |  | YES |  |  | 274 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM |  | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.72 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 7 | SD | [21 empty] | YES |  | YES |  |  | 274 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM |  | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.82 | [6 empty] | MLR | Model 0 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 165 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM |  | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 1 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 165 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM |  | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 2 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 165 | USD | https://www.mdpi.com/1424-8220/18/2/59
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM |  | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 3 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 165 | USD | https://www.mdpi.com/1424-8220/18/2/59
TSI |  | Sharp | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -5.37 | 0.5 |  |  |  |  | 97 | [40 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
Sharp |  |  | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 2.5 | hour | 0.9838 | 91.1 | 196 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 25 | [5 empty] | 26.1 | 26.9 | [8 empty] | YES | 1000 | YES | 0.75 | YES | 40 | YES | 15 |  | 10 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
TSI |  | Sharp | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.64 | 2.64 | 1.14 | 0.41 |  |  |  | 87 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  | Sharp | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.608 | -0.04 | 1.13 |  |  |  |  | 87 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  | Sharp | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.81 | -1.38 | 1.19 |  |  |  |  | 74 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  | Sharp | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.45 | 107 | 3.1 |  |  |  | 13.4 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
TSI |  | Sharp | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -7.3 | 7.7 |  |  |  | 2.6 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
TSI |  | Sharp | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.42 | 122 | 2.8 |  |  |  | 4 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
TZOA |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | TZOA-PM Research Sensors | Y | Y |  | nephelometer | OEM |  | PM | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.4356 | 1.37 | 6.68 | 17 |  |  |  | 47 | [40 empty] | 90 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TZOA |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | TZOA-PM Research Sensors | Y | Y |  | nephelometer | OEM |  | PM | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.518 | 2.16 | 6.75 |  |  |  |  | 47 | [40 empty] | 90 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.86 | -0.004 | 0.56 |  |  |  |  | 82 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.846 | -0.004 | 0.58 |  |  |  |  | 73 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9216 | -0.004 | 0.77 |  |  |  |  | 81 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | Sensoric | NO2 | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9986 | 1.67 | -3.6e-05 |  |  |  |  | linear | Rs=a0+a1*NO2 | O3 |  |  | NO |  |  |  |  | REF | Sensor | [15 empty] | -0.032 ± 0.0032 | -1.22 ± 0.050 | 1.53 ± 0.015 | -2.48 ± 0.049 | 26.8 ± 0.3 | 0.56 ± 0.43 | [7 empty] | 16.7 | ppb.min |  |  | 0-50 | ppb | [10 empty] | YES |  | YES |  |  | 100 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | CO 3E300 | Y | Y |  | electrochemical | OEM | Sensoric | CO | V/ppm | Y | open source | LAB |  |  |  | 1 | hour | 1 | 1.66 | 0.01959 |  |  |  |  | linear | Rs=a0+a1*CO |  |  |  | NO |  |  |  |  | REF | Sensor | [14 empty] | -0.007 ± 111 | 0.50 ± 0.080 |  | -0.74 ± 0.54 | -3.1 ± 5.3 | -25.6 ± 1.4 | 0.26 ± 0.82 | [7 empty] | 0.056 | ppm.min |  |  | 0-500 | ppb | [10 empty] | YES |  | YES |  |  | 100 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | Sensoric | O3 | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9992 | 1.617 | -0.0001858 |  |  |  |  | linear | Rs=a0+a1*O3 |  |  |  | NO | NO2 | CO |  |  | REF | Sensor | [14 empty] | 0.53 ± 0.026 | -0.018 ± 0.0030 | -0.11 ± 0.035 |  | -0.10 ± 0.017 | 6.6 ± 0.90 | 0.058 ± 0.55 | [7 empty] | 3.2 | ppb.min |  |  | 0-1000 | ppb | [10 empty] | YES |  | YES |  |  | 500 | USD | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric | NO | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9947 | 1.32 | 1.001e-05 |  |  |  |  | linear | Rs=a0+a1*NO | O3 |  |  |  | NO2 | CO |  |  | REF | Sensor | [14 empty] | 0.23 ± 0.033 |  | 2.2 ± 0.23 | -0.16 ± 0.033 | 0.52 ± 0.86 | -53.4 ± 0.76 | -0.17 ± 0.052 | [7 empty] | 74.9 | ppb.min |  |  | 0-1000 | ppb | [10 empty] | YES |  | YES |  |  | 120 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 2 | hour | 0.88 | [6 empty] | quad | Rs=(1+0.25*RH2)/(1-RH) |  |  | RH | [5 empty] | BAM | Sensor | linear | 0.75 | -3.22 | 1.22 | [38 empty] | YES |  |  | YES |  |  |  | 0 | EUR | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | RH < 60%
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [11 empty] | RH | [5 empty] | PARTISOL | Sensor | linear | 0.93 | 1.85 | 0.92 | [38 empty] | YES |  |  | YES |  |  |  | 0 | EUR | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [11 empty] | RH | [5 empty] | BAM | Sensor | linear | 0.71 | 1.06 | 0.61 | [38 empty] | YES |  |  | YES |  |  |  | 0 | EUR | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955
Plantower |  |  | Laquai_particle_2017 | 2017 | https://www.researchgate.net/publication/320555036_Particle_Distribution_Dependent_Inaccuracy_of_the_Plantower_PMS5003_low-cost_PM-sensor | PMS5003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [49 empty] | YES | [6 empty] | 15 | EUR | https://www.researchgate.net/publication/320555036_Particle_Distribution_Dependent_Inaccuracy_of_the_Plantower_PMS5003_low-cost_PM-sensor |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | -0.69 | 0.66 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | 0.48 | 0.48 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | -0.53 | 0.65 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Nova Fitness |  |  | Laquai_particle_2017 | 2017 | https://www.researchgate.net/publication/322628807_Development_of_a_Calibration_Methodology_for_the_SDS011_Low-Cost_PM-Sensor_with_respect_to_Professional_Reference_Instrumentation | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour |  | 1.2203 | -0.509 |  |  |  |  | log | RsPM2.5/PM2.5=a1*ln(RSPM10/RSPM2.5)+a0 |  |  | RH | [28 empty] | 16 | SD | [19 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/322628807_Development_of_a_Calibration_Methodology_for_the_SDS011_Low-Cost_PM-Sensor_with_respect_to_Professional_Reference_Instrumentation |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban | 10 |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.61 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban | 10 |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.64 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban | 10 |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.67 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban | 10 |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.773 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.48 |  |  |  |  | 73 | [17 empty] | 54.5 | [24 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.525 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.565 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.67 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.676 |  |  |  |  | 169.3 | [17 empty] | 6.98 | [24 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.71 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.75 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.84 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.3433 |  |  |  |  | 114.6 | [17 empty] | 17.6 | [24 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.353 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.3766 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.556 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3/mV | N | black box | LAB |  |  |  | 15 | min | [17 empty] | Sensor | REF | linear | 0.99 | 266.33 | 9.39 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 48 | hour | [17 empty] | Sensor | GRAVIMETRIC | linear | 0.90253 |  | 1.4969 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3/mV | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.9 | 12 | 2.78 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3/mV | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.96 | -0.05 | 0.92 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | sensor system | OPC-N2 | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | SidePak | linear | 0.497 |  |  | 2.94 | [44 empty] | 0 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | CO-B4 | CO | ppm | N | black box | FIELD | indoor |  |  | 5 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.99 |  |  | 0.01792 | [41 empty] | YES |  |  | 0 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | NO2-B43F | NO2 | ppb | N | black box | FIELD | indoor |  |  | 5 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.93 |  |  | 3.204 | [41 empty] | YES |  |  | 0 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | NO-B4 | NO | ppb | N | black box | FIELD | indoor |  |  | 1 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.97 |  |  | 3.221 | [41 empty] | YES |  |  | 0 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584
UB AirSense-Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | UBAS | Y | N |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.85 | 40 | 1.6 |  |  |  | 4 | [41 empty] | 100 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.008 | 31.649 | 0.0576 |  |  |  |  | 95 |  | 31 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0006 | 19.311 | 0.0457 |  |  |  |  | 96 |  | 20 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0002 | 19.786 | 0.0191 |  |  |  |  | 88 |  | 20 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.814 | 3107 | 0.2113 | 0.2448 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.808 | 18.9 | 0.94 | 26.6 | [16 empty] | 140 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.949 | 2385 | 0.4851 | -1.574 | 0.1029 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.916 | 0.6 | 0.96 | 6.1 | [16 empty] | 22 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.881 | 6142 | -0.1479 | -0.1559 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.839 | -2.4 | 1.04 | 7.7 | [16 empty] | 35 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.914 | 6135 | -0.2636 | -0.2597 | 0.1249 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.84 | -5.9 | 1.1 | 9.8 | [16 empty] | 37 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.903 | 3124 | 0.233 | 0.1435 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.872 | 12 | 0.94 | 21 | [16 empty] | 120 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.95 | 2382 | 0.4505 | -1.039 | 0.08234 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.889 | 2.5 | 0.96 | 7 | [16 empty] | 25 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.872 | 6146 | -0.1409 | -0.05854 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * RH | [8 empty] | REF | Sensor | linear | 0.814 | 1.1 | 1.02 | 8 | [16 empty] | 40 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.902 | 6140 | -0.2222 | -0.2196 | -0.03732 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor,f(T)) + a3 * T | [8 empty] | REF | Sensor | linear | 0.878 | -7.6 | 1.02 | 7.6 | [16 empty] | 22 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.923 | 2379 | 0.4731 | -0.6332 | 0.07075 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.898 | 0.2 | 0.96 | 6.9 | [16 empty] | 27 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.88 | 6145 | -0.149 | -0.1037 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.833 | 0.2 | 0.96 | 7.3 | [16 empty] | 30 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.91 | 6137 | -0.2335 | -0.2364 | -0.1559 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.85 | -4.3 | 1.03 | 8.7 | [16 empty] | 24 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.871 | 3115 | 0.2221 | 0.2002 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.858 | 13.5 | 0.95 | 22.3 | [16 empty] | 120 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.948 | 2386 | 0.4462 | -1.024 | 0.08168 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.93 | 0.3 | 0.98 | 5.4 | [16 empty] | 24 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.871 | 6141 | -0.1544 | -0.2648 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.816 | -1.3 | 0.97 | 7.8 | [16 empty] | 30 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.908 | 6140 | -0.1993 | -0.2273 | -0.04756 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.862 | -6.4 | 1.16 | 9.5 | [16 empty] | 50 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.899 | 3133 | 0.2388 | 0.05537 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.88 | 5.1 | 0.96 | 20.7 | [16 empty] | 120 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.938 | 2384 | 0.468 | -1.228 | 0.08926 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.922 | 0.2 | 0.98 | 5.9 | [16 empty] | 24 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.833 | 6139 | -0.1405 | -0.5694 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.799 | -3.5 | 1.09 | 9.2 | [16 empty] | 55 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.917 | 6138 | -0.2393 | -0.23 | -0.01232 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.894 | -6 | 1.09 | 8.4 | [16 empty] | 30 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.814 | 3107 | 0.2113 | 0.2448 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.864 | 19.5 | 0.94 | 19.7 | [16 empty] | 98 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.949 | 2385 | 0.4851 | -1.574 | 0.1029 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.9332 | 0.3 | 0.98 | 4.9 | [16 empty] | 19 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.881 | 6142 | -0.1479 | -0.1559 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.9215 | -4.3 | 1.09 | 4.9 | [16 empty] | 18 | U at 80 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.914 | 6135 | -0.2636 | -0.2597 | 0.1249 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.875 | -7.2 | 1.13 | 8.7 | [16 empty] | 30 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.903 | 3124 | 0.233 | 0.1435 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.91 | 14.6 | 0.93 | 16.5 | [16 empty] | 102 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.95 | 2382 | 0.4505 | -1.039 | 0.08234 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.889 | 2.4 | 0.97 | 6.3 | [16 empty] | 15 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.872 | 6146 | -0.1409 | -0.0585 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * RH | [8 empty] | REF | Sensor | linear | 0.908 | -0.4 | 1.06 | 5.1 | [16 empty] | 18 | U at 80 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.902 | 6140 | -0.2222 | -0.2196 | -0.03732 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor, f(T)) + a3 * T | [8 empty] | REF | Sensor | linear | 0.91 | -7.6 | 1.01 | 6.4 | [16 empty] | 17 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.923 | 2379 | 0.4731 | -0.6332 | 0.07075 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.911 | -0.6 | 0.97 | 5.6 | [16 empty] | 18 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.88 | 6145 | -0.149 | -0.1037 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.91 | -1.3 | 1.01 | 4.9 | [16 empty] | 9 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.91 | 6137 | -0.2335 | -0.2364 | -0.1559 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.891 | -4.9 | 1.04 | 7.3 | [16 empty] | 22 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.872 | 3115 | 0.2221 | 0.2002 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.89 | 17 | 0.94 | 17.1 | [16 empty] | 75 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.95 | 2386 | 0.4462 | -1.024 | 0.08168 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.948 | -0.1 | 1 | 4.2 | [16 empty] | 17 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.871 | 6141 | -0.1544 | -0.2649 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.898 | -3 | 1.02 | 5.2 | [16 empty] | 11.5 | U at 80 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.908 | 6140 | -0.1993 | -0.2273 | -0.04756 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.901 | -7.3 | 1.18 | 7.8 | [16 empty] | 35 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.899 | 3133 | 0.2388 | 0.05537 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.911 | 6.8 | 0.96 | 15.6 | [16 empty] | 95 | U at 500 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.938 | 2384 | 0.468 | -1.228 | 0.08926 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.941 | 0 | 0.98 | 4.6 | [16 empty] | 13 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.833 | 6139 | -0.1405 | -0.5694 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.889 | -5.5 | 1.15 | 6.2 | [16 empty] | 25 | U at 80 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.917 | 6138 | -0.2393 | -0.23 | -0.01232 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.928 | -6.6 | 1.11 | 6.3 | [16 empty] | 24 | U at 100 | [26 empty] | 1000 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | MLR | Rs = a0 + a1 * NO + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.93 |  | 0.99 | 6.7 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-B4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | ANN | All parameters | [8 empty] | REF | Sensor | linear | 0.97 |  | 0.97 | 4.4 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | MLR | Rs = a0 + a1 * NO2 + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.81 |  | 0.97 | 7.1 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | ANN | All parameters | [8 empty] | REF | Sensor | linear | 0.9 |  | 0.89 | 4.7 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | MLR | Rs = a0 + a1 * O3 + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.8 |  | 0.96 | 10.3 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | ANN | All parameters | [8 empty] | REF | Sensor | linear | 0.86 |  | 0.91 | 8.2 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | MLR | Rs = a0 + a1 * CO + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.81 |  | 0.78 | 0.04 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | JRC_report_Gerboles et al. 2019 | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | ANN | All parameters | [8 empty] | REF | Sensor | linear | 0.8 |  | 0.62 | 0.05 | [44 empty] | 1000 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.788 | 4.8544 | 1.703 |  |  |  |  | 100 | 100 | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7718 | 4.9037 | 2.0427 |  |  |  |  | 100 | 100 | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8113 | 4.375 | 1.9506 |  |  |  |  | 100 | 100 | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.798 | 4.7902 | 1.718 |  |  |  |  | 100 | 100 | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7811 | 4.8434 | 2.0593 |  |  |  |  | 100 | 100 | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8213 | 4.325 | 1.9628 |  |  |  |  | 100 | 100 | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8903 | 3.9835 | 1.9068 | [6 empty] | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8767 | 4.1872 | 2.2491 | [6 empty] | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8991 | 4.033 | 2.0394 | [6 empty] | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6176 | 7.8671 | 0.8689 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6053 | 7.9458 | 1.027 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6911 | 6.4735 | 1.1354 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6948 | 5.9767 | 1.0526 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6916 | 6.1498 | 1.2367 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7619 | 5.2231 | 1.2741 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4064 | 9.638 | 0.6573 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4232 | 9.4925 | 0.8007 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4801 | 8.3687 | 0.882 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7496 | 3.4146 | 1.1942 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7852 | 3.5272 | 1.4267 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8143 | 3.0113 | 1.415 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.523 | 9.3879 | 0.7119 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.5373 | 9.3947 | 0.8577 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.6116 | 8.3596 | 0.939 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.535 | 9.2966 | 0.7225 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.5511 | 9.2799 | 0.8734 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.6243 | 8.2691 | 0.9507 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7928 | 5.3506 | 1.141 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.8103 | 5.5877 | 1.3466 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.8464 | 5.0612 | 1.3403 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5038 | 13.514 | 0.9981 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4827 | 14.702 | 1.1936 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5333 | 11.741 | 1.2428 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2497 | 29.319 | 0.5408 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2254 | 30.739 | 0.6207 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2558 | 27.366 | 0.7088 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2769 | 28.249 | 0.8719 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2756 | 28.708 | 1.0624 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2882 | 26.985 | 1.0764 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6821 | 14.313 | 1.1983 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6716 | 16.516 | 1.4124 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7045 | 13.314 | 1.4761 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4542 | 22.042 | 0.7617 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4502 | 22.757 | 0.9234 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4713 | 21.075 | 0.9467 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4615 | 22.376 | 0.8111 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4514 | 23.045 | 0.98 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4858 | 210.23 | 1.0069 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7066 | 11.438 | 1.0602 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7076 | 13.102 | 1.2603 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | open source | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7444 | 10.193 | 1.3189 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7041 | 2.1213 | 0.7802 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6825 | 1.7206 | 0.799 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6854 | 1.499 | 0.8223 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7348 | 1.9365 | 0.7988 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7119 | 1.523 | 0.8178 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7166 | 1.2854 | 0.8434 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9058 | 1.5949 | 0.8316 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8705 | 1.2561 | 0.8386 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.881 | 0.9495 | 0.8746 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5649 | 4.9727 | 0.5269 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5496 | 5.1093 | 0.5147 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5552 | 4.9914 | 0.5287 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.612 | 4.7844 | 0.5383 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5836 | 4.9388 | 0.5247 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5901 | 4.8162 | 0.5395 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7725 | 4.1765 | 0.5413 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7451 | 4.9157 | 0.5226 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7416 | 4.88 | 0.534 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5485 | 5.6826 | 0.5942 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5184 | 5.9531 | 0.5724 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5135 | 0.59431 | 0.5811 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.5714 | 6.2735 | 0.5421 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.534 | 6.6015 | 0.515 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.5228 | 6.6436 | 0.5216 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0507 | 26.368 | 0.5531 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0401 | 27.43 | 0.4931 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0388 | 27.487 | 0.4914 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0602 | 26.139 | 0.5639 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0475 | 27.255 | 0.5016 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0459 | 27.306 | 0.5 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0655 | 27.323 | 0.4978 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0569 | 27.909 | 0.4621 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0488 | 27.445 | 0.4353 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0672 | 32.293 | 0.6137 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0529 | 36.496 | 0.5457 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0491 | 36.752 | 0.533 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.1033 | 34.746 | 0.643 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0896 | 35.494 | 0.5971 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0811 | 35.901 | 0.5777 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9977 | -0.6049 | 1.0178 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9926 | -0.8315 | 0.9875 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9893 | 0.0249 | 1.0005 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9988 | -0.6477 | 1.0187 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9955 | -0.976 | 0.9902 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9972 | -0.603 | 1.0128 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9911 | 0.8018 | 0.8018 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9917 | 0.9733 | 0.1887 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9923 | 0.9911 | 0.4417 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9514 | -1.4532 | 1.438 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9684 | -1.2815 | 1.2343 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9492 | -22.029 | 1.5153 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9635 | -6.6195 | 1.2611 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.961 | -1.4771 | 1.4403 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9761 | -1.0974 | 1.2298 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9612 | -22.026 | 1.5157 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9678 | -6.4263 | 1.2545 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.957 | 1.1417 | -0.8297 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9653 | 1.2006 | -0.4091 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9612 | 1.4787 | -20.639 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9641 | 1.2035 | -4.7653 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppb | Y | black box | LAB |  |  | 60 | 1 | min | 0.9909 | -11.51 | 1.3291 |  |  |  |  | linear | [9 empty] | Sensor | REF | linear | [48 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/aeroqual-s500-o3---lab-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8386 | 0.0018 | 0.98029 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/aeroqual-s500-o3---lab-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8554 | 0.0013 | 1.0061 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8606 | 6e-05 | 1.0064 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8571 | 0.0014 | 0.991 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8743 | 0.001 | 1.0157 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8792 | -0.0003 | 1.0145 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8686 | 0.0013 | 0.9939 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8922 | 0.0008 | 1.0202 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8962 | -0.0004 | 1.0185 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8604 | -50.299 | 4.9091 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8374 | -34.185 | 3.1878 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8585 | -49.063 | 4.569 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.868 | 52.791 | 5.0839 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.848 | -35.783 | 3.2562 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8658 | -51.546 | 4.7615 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8731 | -45.053 | 4.3547 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.838 | -26.041 | 2.588 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8731 | -40.097 | 3.737 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0 | 0 | 0.1253 |  |  |  |  | 100 |  | 90 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1442 | 30.209 | -0.081 |  |  |  |  | 100 |  | 160 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1737 | 31.343 | -0.0786 |  |  |  |  | 100 |  | 190 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8656 | -0.0595 | 0.8401 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.868 | -0.0514 | 0.8563 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8755 | 0.0746 | 0.8294 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8423 | 3.4367 | 0.5649 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8543 | 3.513 | 0.589 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.841 | 3.5134 | 0.551 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1235 | 24.687 | 0.5098 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1304 | 26.38 | 0.551 |  |  |  |  | 99.8 |  | 20 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1283 | 23.812 | 0.517 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8686 | -0.1195 | 0.8454 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8736 | -0.1482 | 0.8651 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8787 | 0.0161 | 0.8346 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8563 | 3.4088 | 0.5661 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8703 | 3.467 | 0.5914 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8544 | 3.4895 | 0.5522 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8693 | 0.3665 | 0.8092 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8313 | 0.6412 | 0.8136 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9624 | -0.4642 | 1.0993 | [8 empty] | > 92% (spike) | [36 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9566 | -28.017 | 1.6529 | [45 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8932 | -22.771 | 1.5406 | [45 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.0611 | 10.819 | 0.1644 |  |  |  |  | 86 | 97 | 38 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.2363 | 5.7843 | 0.1953 |  |  |  |  | 86 | 97 | 60 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.0061 | 14.377 | 0.0396 |  |  |  |  | 86 | 97 | 60 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1453 | 8.4929 | 0.1627 |  |  |  |  | 86 | 97 | 58 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4326 | 15.048 | 0.6426 |  |  |  |  | 95 |  | 10 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7175 | 13.507 | 0.2062 |  |  |  |  | 96 |  | 30 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7156 | 5.8943 | 0.6287 |  |  |  |  | 88 |  | 18 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | N | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.826 | 12.306 | 0.0379 | -0.977 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.716 | 9.6771 | 0.0195 | -0.004 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.7676 | 9.6273 | 0.0186 | -0.0701 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.8301 | 12.322 | 0.0381 | -0.9875 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7221 | 9.2623 | 0.0194 | 0.0049 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7726 | 9.2604 | 0.0186 | -0.00621 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.7229 | 3.0643 | 0.0182 | 0.1326 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.6343 | 3.3406 | -0.0124 | 1.6537 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.6798 | -1.8537 | -0.0069 | 1.3516 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt410 | electrochemical | sensor system |  | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4868 | 0.0795 | -0.2051 |  |  |  |  | 85 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt411 | electrochemical | sensor system |  | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4001 | 0.0744 | -0.2761 |  |  |  |  | 85 |  | 0.13 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt412 | electrochemical | sensor system |  | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.5833 | 0.084 | -0.1855 |  |  |  |  | 85 |  | 0.24 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt413 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.6653 | 9.899 | 0.7531 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt414 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7077 | 6.6974 | 0.8982 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt415 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8118 | 2.5717 | 0.9888 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt416 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6679 | 9.014 | 0.7637 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt417 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.7145 | 5.6011 | 0.9164 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt418 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8148 | 2.1074 | 0.9954 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt419 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.4434 | 22.405 | 0.3863 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6161 | 17.796 | 0.5344 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT410 v.1.15 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt421 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6853 | 14.53 | 0.623 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
FIS |  |  | Spinelle et al. 2016 | 2016 | https://www.aidic.it/cet/16/54/054.pdf | FIS SP-61 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | LAB |  |  | 89 | 1 | hour | 0.9991 | 10.8 | 2.86 | 0.00235 | -0.00011 |  |  | MLR | Rs=a0+a1*O3+a2*O3²+a3*O33 | [8 empty] | REF | Sensor | linear | [48 empty] | 50 | USD | https://www.aidic.it/cet/16/54/054.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8859 | 0.2977 | 0.8199 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.7575 | 4.311 | 0.518 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.7466 | 4.6687 | 0.5327 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.7915 | 3.9654 | 0.5335 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.5929 | 5.3928 | 0.5294 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6356 | 5.5364 | 0.5768 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.5769 | 5.3565 | 0.5157 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6246 | 5.5996 | 0.519 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6447 | 6.0854 | 0.5413 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6598 | 5.0593 | 0.5428 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1134 | 32.337 | 0.53 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1156 | 33.645 | 0.5179 |  |  |  |  | 99.8 |  | 20 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1121 | 31.614 | 0.5258 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR P10 V2 | Y | Y | https://www.kunak.es/en/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.54 | -0.33 | 0.66 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR P10 V2 | Y | Y | https://www.kunak.es/en/ | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.82 | -4 | 0.76 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 | -3.1 | 0.95 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 | 2.1 | 0.94 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9 | 0.29 | 0.95 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.96 | 1.7 | 0.94 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.93 | 0.93 | 0.96 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.93 | 3.2 | 0.92 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.43 | 1.26 | 0.55 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.27 | 243 | 0.41 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2017 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.88 | 3.251 | 0.9407 | 9.29 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2018 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | O3 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.91 | 0.2077 | 1.0039 | 6.52 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2018 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | NO | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.95 | 0.7292 | 0.9974 | 10.05 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 16.82 | 1.01 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 2.53 | 0.96 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.74 | 2.35 | 1 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.83 | 7.76 | 0.97 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 2.53 | 0.96 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.78 | -0.94 | 0.94 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.71 | 3.23 | 0.93 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 | -0.27 | 0.9 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.84 | 2.48 | 0.86 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.88 | 5.83 | 0.81 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.85 | 3.35 | 0.82 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | electrochemical | sensor system | A3CO-Citytech | CO | mg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.8 |  | 1.01 | 0.066 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | electrochemical | sensor system | KWJ-Engineering-SNL-NO2 | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.71 |  | 1.01 | 4.76 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | electrochemical | sensor system | SGX-SensorTech-MiCS_OZ 47 | O3 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.75 |  | 1.01 | 5.08 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.9 |  |  | 4.12 | [44 empty] | 3538 | EUR | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf |  | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-OX-B431 | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.97 |  |  | 2.05 | [44 empty] | 3538 | EUR | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO)*T | [8 empty] | REF | Sensor | linear | 0.9 |  |  | 4.6 | 3.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.1 | 2.3 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.95 |  |  | 3.3 | 2.1 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO2 = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO2)*T | [8 empty] | REF | Sensor | linear | 0.6 |  |  | 5.6 | 4.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.74 |  |  | 4.6 | 3.5 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.79 |  |  | 4.1 | 3.2 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO)*T | [8 empty] | REF | Sensor | linear | 0.8 |  |  | 8.3 | 5.8 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 7.1 | 4.2 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 7.1 | 4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO2 = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO2)*T | [8 empty] | REF | Sensor | linear | 0.76 |  |  | 5.5 | 4.4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.83 |  |  | 4.6 | 3.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 4.4 | 3.4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.71 | 0.965 | 0.891 | 4.21 | 2.97 | 0.04 | [14 empty] | 4.21 | SD | 9.64 |  |  | 98.16 | % | [21 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.68 | 2.97 | 0.816 | 4.49 | 3.18 | -1.25 | [14 empty] | 4.31 | SD | 9.64 |  |  | 86.82 | % | [21 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.55 | 0.867 | 0.708 | 5.45 | 3.84 | 1.87 | [14 empty] | 5.12 | SD | 9.64 |  |  | 80.76 | % |  |  |  |  | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.75 | 2.25 | 0.77 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.7 | 2.61 | 0.736 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.6 | 3.57 | 0.635 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.8 | 1.67 | 0.824 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.79 | 1.76 | 0.82 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | open source | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.76 | 2.21 | 0.775 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK

## references
reference
kumar_rise_2015
cen_ambient_2012_CO
cen_ambient_2012_NO2
cen_ambient_2012_O3
cen_ambient_2012_SO2
cen_ambient_2014_PM
lewis_validate_2016
spinelle_evaluation_2017
key-vocs_metrology_2017
Spinelle_et_al_2013
Spinelle_et_al_2013(a)
lewis_low-cost_2018
barrett_coefficient_1974
the_world_air_quality_index_sensing_2019
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(b)
Williams et al. 2014(b)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Borrego et al. 2016
US-EPA_2015
Williams et al. 2014(b)
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Williams et al. 2014(b)
Williams et al. 2014(b)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Sousan et al 2016
Northcross et al. 2013
Northcross et al. 2013
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Steinle et al. 2015
Steinle et al. 2015
Han et al,.2016
Han et al,.2016
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Dacunto et al. 2015
Northcross et al. 2013
Williams et al. 2014(b)
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Borghi et al. 2018
Borghi et al. 2018
Borghi et al. 2018
Borghi et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Sun et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Cavaliere et at. 2018
Cavaliere et at. 2018
Cavaliere et at. 2018
Cavaliere et at. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Spinelle et al. 2015 (a)
Spinelle et al. 2015 (a)
Spinelle et al. 2015 (a)
Spinelle et al. 2015 (a)
Spinelle et al. 2015 (a)
Spinelle et al. 2015 (a)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
Spinelle et at. 2015 Part B
US-EPA_2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
AQ-SPEC
Zimmerman_et_al_2018
Zimmerman_et_al_2018
Zimmerman_et_al_2018
Zimmerman_et_al_2018
Zimmerman_et_al_2018
Zimmerman_et_al_2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Vaughn D et al. 2010 (Sonoma Technology)
Vaughn D et al. 2010 (Sonoma Technology)
Vaughn D et al. 2010 (Sonoma Technology)
Vaughn D et al. 2010 (Sonoma Technology)
Sun et al. 2017
Lin et al. 2015
Karagulian et al. 2012
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Mead et al. 2013
Mead et al. 2013
Wang et al. 2015
Williams et al. 2014(b)
Williams et al. 2014(b)
Cross et al. 2017
Cross et al. 2017
Cross et al. 2017
Cross et al. 2017
Borrego et al. 2016
Borrego et al. 2016
AQ-SPEC
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Feinberg et al. 2018
Feinberg et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Wang et al. 2015
Alvarado et al. 2015
Williams et al. 2014(b)
Olivares and Edwards, 2015
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Sousan et al 2016
Sousan et al 2016
Alvarado et al. 2015
Alvarado et al. 2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Williams et al. 2014(b)
Williams et al. 2014(b)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Austin et al. 2015
Gao et al. 2015
Gao et al. 2015
Kelly et al 2017
Manikonda et al. 2016
Zikova et al. 2016
Zikova et al. 2016
Zikova et al. 2016
Zikova et al. 2016
Zikova et al. 2016
Zikova et al. 2016
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Kelly et al 2017
Kelly et al 2017
Kelly et al 2017
Kelly et al 2017
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
AQ-SPEC
Mead et al. 2013
Mead et al. 2013
Mead et al. 2013
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
AQ-SPEC
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
AQ-SPEC
Wang et al. 2015
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Gerboles et al. 2015(a)
Gerboles et al. 2015(a)
Gerboles et al. 2015(a)
Gerboles et al. 2015(a)
Chakrabarti et al. 2004
Chakrabarti et al. 2004
Chakrabarti et al. 2004
Laquai_particle_2017
Budde et al. 2018
Budde et al. 2018
Budde et al. 2018
Laquai_particle_2017
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Pillarisetti et al 2017
Pillarisetti et al 2017
Pillarisetti et al 2017
Pillarisetti et al 2017
Gillooly et al. 2019
Gillooly et al. 2019
Gillooly et al. 2019
Gillooly et al. 2019
Manikonda et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
JRC_report_Gerboles et al. 2019
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Spinelle et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Bettair 2017
Bettair 2017
Bettair 2017
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Marjovi et al. 2017
Marjovi et al. 2017
Marjovi et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019

## Notes
NOTES
1 Sensor system is intented to be a sensor together with a DAQ (data acquisition) and data transfer; Sensor is intented to be the measuring device only.
3Parameter used for the MLR in addtion to the reference
2Rs is refered to the signal response of the sensor, NO2, O3, NO, CO, T, RH are the intereferents
OEM = Original Equipment Manufacturer
Response time | t90, time necessary to reach 90 % of the final value
Resolution time | Lowest time for acqusition of a new data
Bias
Range | tange of values of the measured pollutant: minimum - maximum
MOdel type for calibration | type of calibration model equation applied to raw data of the sensor, when empty raw data are use, MLR: multi linear regression, RF, random forest, ANN: artifical neural networ, log: logarithmic, exp: exponential

## Sheet2
Honeywell HPM | HPM series | optical | sensor | PM2.5 | N | black box |  |  | up to 95% |  |  | 15ug/m3 (o-100 ug/m3); 15% 100-1000 ug/m3 | 0-1000 ug/m3 |  |  | https://sensing.honeywell.com/sensors/particle-sensors | 22 | EUR
Winsen ZH03A | ZH03 - Laser Dust Module | optical | sensor | PM2.5 | N | black box |  |  |  | PM2.5=1000x(TH)/(TH+TL) (low & high lev. Time) | [5 empty] | https://www.winsen-sensor.com/d/files/air-quality/zh03-series-laser-dust-module-v2_0.pdf | 72 | EUR
PLANTOWER PMS A003 | [15 empty] | http://www.plantower.com/content/?125.html
PLANTOWER PMS 1003
Shinyei | PM Sensor Evaluation Kit | optical | sensor system | PM2.5 | N | black box | 0.92 | [8 empty] | https://www.shinyei.co.jp/stc/eng/optical/main_pmmonitor.html | 5200 | EUR
Dylos | DC 1700 | optical | sensor system | PM2.5 | N | black box | 0.917 | [8 empty] | http://www.dylosproducts.com/dc1700.html | 425 | EUR
PMX | Custom PMX | optical | sensor system | PM2.5 | N | open source | 0.97 | 0.1905 |  |  | 30.8 ug/m3 (std) |  |  |  |  | Park_et_al_2013
Air nut sensor |  | optical | sensor system | PM2.5 | N | black box | 0.81 | [6 empty] | 20-23 | 20-23 | file:///C:/Users/karagfe/Downloads/SDS011_laser_PM2.5_sensor_specification-V1.3.pdf | 200 | EUR
Nova Fitness Co. Ltd. | SDS011 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m3 |  | 0-1000 ug/m3 |  |  | https://pan.baidu.com/s/1skCk8AL | 20 | EUR
Nova Fitness Co. Ltd. | SDS018 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m4 |  | 0-1000 ug/m4 |  |  | https://ecksteinimg.de/Datasheet/SDS018%20Laser%20PM2.5%20Product%20Spec%20V1.5.pdf | 20 | EUR
Nova Fitness Co. Ltd. | SDS021 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m5 |  | 0-1000 ug/m5 |  |  | https://cdn.sparkfun.com/assets/parts/1/2/2/7/5/SDS021_laser_PM2.5_sensor_specification-V1.0.pdf | 20 | EUR
Nova Fitness Co. Ltd. | SDS019 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m6 |  | 0-1000 ug/m6 |  |  | https://pan.baidu.com/s/1sl7Hj0h | 21 | EUR
Met One | ES-642 | optical | sensor | PM2.5 | N | open source | 0.92
Winsen | ZE03-NO2 | electrochemical | sensor | NO2 | N | open source | [9 empty] | https://aqicn.org/air/view/sensor/spec/no2.winsen-ZE03.pdf | 90 | EUR
Winsen | ZE03-CO | electrochemical | sensor | CO | N | open source | [9 empty] | https://aqicn.org/air/view/sensor/spec/no2.winsen-ZE03.pdf | 90 | EUR
Alphasense | NO2-A1
https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance
EPA reports
https://www.epa.gov/air-sensor-toolbox/air-sensor-toolbox-resources-and-funding#TER
https://www.epa.gov/sites/production/files/2018-03/documents/final_em-3_master_slide_set.pdf