## DataBase
Manufacturer | User - Buyer | Project | reference | year_ref | link_reference | Model | Commercial | living | Manufacturer Web link | type(opc, nephelometer, electrochemical, metal oxide) | OEM*/ Sensor system1 | OEMFirst provider | Pollutant | Sensor result unit | Availability of raw data | Open source correction or black box? | Field / Lab | site | Response Time (s) | Time resolution (s) | Averaging time | units (avg_time) | R2_calib | intercept(a0) | a1(slope) | a2 | a3 | a4 | a5 | model type calibration | calibration type equation2 | Int_O3 | Int_T | Int_RH | Int_NO | Int_NO2 | Int_CO | Int_CO2 | Int_day | X | Y | model type regression (comparison) | r2 | intercept | slope | RMSE | MAE | BIAS | nRMSE (%) | min data recovery(%) | max data recovery(%) | Intra-model variability mean Field Value (µg/m3) | T (°C) | RH (%) | % failure | NO2, Sensitivity Coeff (ppb/ppb or ppm) | NO, Sensitivity Coeff (ppb/ppb or ppm) | CO, Sensitivity Coeff (ppb/ppm) | O3, Sensitivity Coeff (ppb/ppb or ppm) | SO2, Sensitivity Coeff (ppb/ppb or ppm) | Temp, Sensitivity Coeff (ppb/°C) | RH, Sensitivity Coeff (ppb/%) | U | type of U | RSD (100-1000 µg/m3) (%) | RSD (< 100 µg/m3) (%) | min Accuracy | max Accuracy | units Accuracy | LoD | units (LoD) | LoD (100-1000 µg/m3) µg/m3 | LoD (< 100 µg/m3) µg/m3 | RANGE | units RANGE | precison | units (precision) | repteability (at 100 ppb) | Reproducibility between sensors (R2) | Dss-drift | Dls-drift (ppb/day) | Dep. Part. Composition | impact Part. Compositions (%) | Dep. Part. Size | impact part. Size (%) | Dep. RH | impact RH (%) | Dep. T | impact T (%) | Long Term Performance | Price | currency | link_reference | comments | operator (added by)
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | [48 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.402 | 15.173 | -0.1134 |  |  |  |  | 90 |  | 102 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3974 | 14.893 | 0.0002 |  |  |  |  | 96 |  | 100 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N |  | MOs | sensor system | SGX Sensortech MICS-4514 | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3381 | 14.002 | 0.0001 |  |  |  |  | 95 |  | 120 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8604 | -50.299 | 4.9091 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8374 | -34.185 | 3.1878 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8585 | -49.063 | 4.569 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.868 | 52.791 | 5.0839 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.848 | -35.783 | 3.2562 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8658 | -51.546 | 4.7615 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8731 | -45.053 | 4.3547 |  |  |  |  | 97 |  | 50 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.838 | -26.041 | 2.588 |  |  |  |  | 40 |  | 28 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf | Egg v.1 | Y | N | https://shop.wickeddevice.com/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8731 | -40.097 | 3.737 |  |  |  |  | 92 |  | 21 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v1---field-evaluation.pdf |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9977 | -0.6049 | 1.0178 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9926 | -0.8315 | 0.9875 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9893 | 0.0249 | 1.0005 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9988 | -0.6477 | 1.0187 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9955 | -0.976 | 0.9902 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9972 | -0.603 | 1.0128 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9911 | 0.8018 | 0.8018 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9917 | 0.9733 | 0.1887 |  |  |  |  | 90 |  | 40 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
2B Technologies |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 | 2B Tech. (POM) | Y | Y | https://twobtech.com/products.html | UV | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 10 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9923 | 0.9911 | 0.4417 |  |  |  |  | 90 |  | 39 | [38 empty] | 4500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/2b-technologies-pom---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4158 | 223.36 | 0.3731 |  |  |  |  | 93 |  | 157 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.8087 | 196.16 | 0.5864 |  |  |  |  | 100 |  | 157 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.7928 | 206.39 | 0.6051 |  |  |  |  | 90 |  | 117 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0439 | 7.1815 | 0.1555 |  |  |  |  | 93 |  | 10 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4568 | 2.673 | 1.2113 |  |  |  |  | 100 |  | 5 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense - NO2-B43F | NO2 | ppb | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.1103 | 3.6067 | 0.3914 |  |  |  |  | 90 |  | 10 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0032 | 3.6916 | 0.0045 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.4393 | 1.9207 | 0.7121 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
AQMesh |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | AlphaSense | NO | ppm | N | black box | FIELD | rural |  | 900 | 15 | min | [17 empty] | Sensor | FRM | linear | 0.0075 | 3.4064 | 0.014 | [45 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/aqmesh-(v-4-0)---field-evaluation.pdf |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | 0.814 | 3.977 | -8e-12 | 5e-05 |  |  |  | quad | PM=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | ` |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM |  | 0.83 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 1 | hour | 0.632 | 4.1744 | -1e-11 | 4e-05 |  |  |  | quad | PM=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | BAM | [8 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM |  | 0.89 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM |  | 0.81 | [6 empty] | 100 |  | 330000 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/dylos-dc1100---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Dylos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/dylos---lab-evaluation.pdf?sfvrsn=2 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/L | Y | open source | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.89 | -23404 | 0.93 |  |  |  |  | 97 | 100 |  | 20 | 40 | [12 empty] | -4147 | -117 | [22 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/dylos---lab-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0022 | 9.2911 | 0.0357 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0166 | 7.8293 | 0.0982 |  |  |  |  | 100 |  | 20 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0203 | 7.752 | 0.1013 |  |  |  |  | 100 |  | 20 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1463 | 19.254 | 0.3471 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1479 | 1.7333 | 0.5978 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0026 | 32.344 | 0.0607 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.016 | 28.119 | 0.09 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.1714 | 0.6979 | 0.6428 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0023 | 32.472 | 0.0554 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0142 | 38.853 | 0.00958 |  |  |  |  | 100 |  | 60 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.1764 | 5.954 | 0.728 |  |  |  |  | 100 |  | 53 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0091 | 42.256 | 0.1243 |  |  |  |  | 100 |  | 18 | [38 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9624 | -0.4642 | 1.0993 | [8 empty] | > 92% (spike) | [36 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9566 | -28.017 | 1.6529 | [45 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Perkin Elmer |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 | ELM | Y | N |  | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8932 | -22.771 | 1.5406 | [45 empty] | 5200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/perkin-elmer-elm---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.52 | 10.394 | 0.461 |  |  |  |  | 77 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.563 | 11.136 | 0.619 |  |  |  |  | 77 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.54 | 11.04 | 0.392 |  |  |  |  | 77 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.533 | 12.385 | 0.5068 |  |  |  |  | 77 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6609 | 9.12 | 0.52 |  |  |  |  | 90 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.674 | 10.087 | 0.7 |  |  |  |  | 90 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.657 | 9.848 | 0.466 |  |  |  |  | 90 |  | 10.5 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 | Met One (NM) | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 900 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 10.966 | 0.605 |  |  |  |  | 90 |  | 5.2 | [38 empty] | 1900 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-neighborhood-monitor---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8085 | 5.5525 | 0.487 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8003 | 8.6212 | 0.5092 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8684 | 4.3356 | 0.5592 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8205 | 5.4635 | 0.4924 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8102 | 8.2603 | 0.5227 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8762 | 4.3305 | 0.5548 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8436 | 5.3419 | 0.5 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7289 | 6.7195 | 0.8546 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8967 | 4.0953 | 0.566 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7809 | 6.8335 | 0.5547 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6743 | 9.7879 | 0.5939 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.801 | 5.7815 | 0.6133 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8956 | 6.4465 | 0.5919 |  |  |  |  | 30 | 80 | 13.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7733 | 9.2457 | 0.6421 |  |  |  |  | 30 | 80 | 12.5 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9105 | 5.1977 | 0.6545 |  |  |  |  | 30 | 80 | 14 | [38 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---field-evaluation.pdf |  | JRC-FK
RTI |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---lab-evaluation.pdf | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | 9.41 | 0.31 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -27- (-101.5) over 0-240 ug/m3 | [23 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/rti-micropem---lab-evaluation.pdf |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.634 | 4.037 | 1.34 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.705 | 5.217 | 5.681 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.818 | 3.59 | 5.823 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.654 | 6.11 | 0.713 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.706 | 5.179 | 5.683 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.822 | 3.535 | 53838 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.489 | 10.873 | 0.48 |  |  |  |  | 100 | 100 | 39 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.545 | 4.988 | 5.162 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.808 | 2.754 | 5.951 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.672 | 3.745 | 1.419 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.451 | -25.326 | 16.939 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.567 | -34.899 | 18.658 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [17 empty] | 0.01 | µg/m3 | [14 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.698 | 5.652 | 0.7534 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.465 | -25.426 | 16.915 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.597 | -35.437 | 18.763 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [7 empty] | linear | PM=a1*Rs+a0 | [8 empty] | Sensor | GRIMM | linear | 0.524 | 9.909 | 0.505 |  |  |  |  | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.579 | -27.738 | 17.412 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.853 | -43.518 | 20.86 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.583 | 3.664 | 1.451 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.7664 | -42.362 | 20.108 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.756 | 4.809 | 0.83 |  |  |  |  | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.4175 | 4.39 | 0.647 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.527 | 3.647 | 0.711 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.434 | 4.092 | 0.616 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.658 | 3.343 | 0.7368 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.673 | 3.319 | 0.86 |  |  |  |  | 100 | 100 | 2.2 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.38 | 8.882 | 0.665 |  |  |  |  | 100 | 100 | 3 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.395 | 8.804 | 0.594 |  |  |  |  | 100 | 100 | 3.8 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.919 | 1.874 | 0.857 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.822 | 1.979 | 0.8245 |  |  |  |  | exp | PM=a0*a1*Rs | [8 empty] | Sensor | BAM | [8 empty] | 100 | 100 | 11 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.904 | 2.617 | 0.906 |  |  |  |  | 100 | 100 | 30 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.826 | 5.982 | 0.952 |  |  |  |  | 100 | 100 | 37 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.763 | 6.5 | 0.796 |  |  |  |  | 100 | 100 | 39 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.659 | -2.59 | 8.88 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | 0.794 | -5.29 | 9.182 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.663 | -2.7402 | 8.934 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -7.8 | 0.64 |  |  |  |  | 42 | 100 |  | 20 | 40 | [12 empty] | 0.3 | 43.8 | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -8.98 | 0.56 |  |  |  |  | 42.4 | 100 |  | 20 | 40 | [12 empty] | -33.2 | 1780% | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | BAM | linear | 0.99 | -1.72 | 0.97 |  |  |  |  | 9.5 | 96.8 |  | 20 | 40 | [12 empty] | 85.1 | 9590% | % | [16 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/alphasense-opc-n2---lab-evaluation.pdf?sfvrsn=8 | low precison: T=5° and 15% RH | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.027 | 4.5149 | 0.0352 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1057 | 2.864 | 0.0829 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0189 | 4.654 | 0.0284 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0985 | 8.3815 | 0.1097 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.2159 | 5.2826 | 0.1931 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.086 | 8.522 | 0.0988 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.25 | 5.52 | 0.06 |  |  |  |  | 88.7 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.21 | 5.98 | 0.05 |  |  |  |  | 95.4 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.33 | 4.15 | 0.06 |  |  |  |  | 92.4 | [40 empty] | 200 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.67 | -323 | 273 | 0.06 |  |  |  | 68 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7056 | -1234 | 278 |  |  |  |  | 68 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | hppcf | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.67 | -352 | 322 |  |  |  |  | 29 | [40 empty] | 200 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.826 | -9.522 | 1.849 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089 |  | JRC-FK
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | min | [17 empty] | Sensor | REF | linear | 0.861 | 6.632 | 0.575 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089 |  | JRC-FK
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.968 | -0.801 | 0.629 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089 |  | JRC-FK
HabitatMap |  | AirCasting | Borghi et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/3089 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 8 | hour | [17 empty] | Sensor | REF | linear | 0.889 | -2.67 | 1.808 | [45 empty] | 200 | USD | https://www.mdpi.com/1424-8220/18/9/3089 |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.94 | 2.56 | 888.26 | [6 empty] | 18.3 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.94 | 2.58 | 811.52 | [6 empty] | 17.6 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Aeroqual-SM50 | Y | Y |  | MOs | sensor system | Aeroqual-SM50 | O3 | ppm | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.82 | 2.91 | 955.26 | [6 empty] | 18.2 | [35 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.42 | 5.24 | 0.002 | [6 empty] | 11.2 | [12 empty] | 0.62 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.5 | 12.2 | 0.00203 | -0.0693 | -0.016 |  |  | MLR | NO2=a0+a1*Rs+a2*T+a3*RH |  | T | RH | [5 empty] | Sensor | BAM | linear | 0.43 | 6.01 | 0.0018 | [6 empty] | 11.48 | [12 empty] | 0.58 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AirBeam | Y | Y |  | nephelometer | sensor system | HabitatMap-AirBeam - Shineyi PPD60PV | PM2.5 | hppcf | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.43 | 6.4 | 0.0017 | [6 empty] | 11.45 | [12 empty] | 0.61 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AlphaSense | CO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.75 | -0.0709 | 0.000888 | 0.002252 |  |  |  | MLR | CO=a0+a1*Rs+a2*day | [7 empty] | day | Sensor | FEM | linear | 0.63 | 0.0695 | 0.000809 | [6 empty] | 326.2 | [12 empty] | 0.7 | [25 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AlphaSense | CO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.68 | 0.0512 | 0.000799 | [6 empty] | 326.2 | [12 empty] | 0.68 | [25 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AlphaSense | NO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.75 | [6 empty] | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day | [8 empty] | Sensor | FEM | linear | 0.77 | 1.06 | 0.892 | [6 empty] | 18.3 | [12 empty] | 1.84 | [22 empty] | YES |  |  | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | AQMesh v.3.0 | Y | N |  | electrochemical | sensor system | AlphaSense | NO | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.87 | -0.765 | 0.883 | [6 empty] | 18.5 | [12 empty] | 1.76 | [22 empty] | YES |  |  | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | sensor system | Envea - CairClip O3/NO2 - Membrapor O3/M-5 | O3 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.709 | 26.6 | -0.175 |  |  |  |  | MLR | O3 = a0 + a1*Rs - a2*RH |  |  | RH | [5 empty] | Sensor | REF | linear | 0.88 | 12.451 | 0.88 | [6 empty] | 18.6 | [35 empty] | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | sensor system | Envea - CairClip O3/NO2 - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.81 | 26.2 | 1.18 | -0.483 | -0.129 |  |  | MLR | Rs(NO2)=Rs-O3; NO2=a0+a1*Rs+a2*T+a3*RH |  | T | RH | [5 empty] | Sensor | FEM | linear | 0.57 | 11.391 | 0.955 | [6 empty] | 11.6 | [12 empty] | 40 | [20 empty] | YES |  | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | CairClip O3/NO2 | Y | Y |  | electrochemical | sensor system | Envea - CairClip O3/NO2 - Membrapor O3/M-5 | O3 | ppb | Y | black box | FIELD | suburban |  |  | 1 | hour | [17 empty] | Sensor | REF | linear | 0.68 | 12.203 | 0.68 | [6 empty] | 16.7 | [35 empty] | YES |  |  | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 | Y | Y |  | OPC | sensor system | Dylos DC1100 | PM2.5-0.5 | #/(ft3/100) | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.33 | 8.08 | 0.0086 | [6 empty] | 11.71 | [12 empty] | 0.83 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 PRO | Y | Y |  | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/(ft3/100) | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.6 | 14.8 | 0.00164 | -0.0751 | -0.035 |  |  | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day |  |  | RH |  |  |  |  | day | Sensor | BAM | linear | 0.45 | 7.2 | 0.0015 | [6 empty] | 11.67 | [12 empty] | 0.77 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | Dylos DC1100 PRO | Y | Y |  | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/(ft3/100) | Y | black box | FIELD | suburban |  |  | 12 | hour | [17 empty] | Sensor | BAM | linear | 0.4 | 7.78 | 0.0013 | [6 empty] | 11.64 | [12 empty] | 0.85 | [22 empty] | YES |  |  | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
CAIRSENSE |  | CAIRSENSE | Jiao et al. 2016 | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system | Shinyei PMS-SYS-1 | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.42 | 11.3 | 0.78 | -0.04 | -0.0112 |  |  | MLR | PM2.5=a0+a1*Rs+a2*RH+a3*day |  |  | RH |  |  |  |  | day | Sensor | FEM | linear | 0.36 | 0.748 | 0.72 | [6 empty] | 12.18 | [12 empty] | 0.69 | [25 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Jiao%20et%20al_2016_Community%20Air%20Sensor%20Network%20(CAIRSENSE)%20project.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0576 | 13 | 0.76 | 37 |  |  |  | 93 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.16 | 15 | 0.74 |  |  |  |  | 96 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Speck | Y | Y | https://www.elecrow.com/download/DSM501.pdf | nephelometer | sensor system | SAMYYoung S&C - DSM501A | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.1225 | 10 | 0.62 |  |  |  |  | 61 | [40 empty] | 150 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | US-EPA_2015 | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.039 | 2.76 | 0.146 | [45 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | temperature related correction | JRC-FK
Airviz |  | Carnegie Mellon | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts | N | black box | FIELD | rural |  | 1 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0342 | 119.11 | -6.5434 | [8 empty] | > 95% (spike) | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Airviz |  | Carnegie Mellon | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 1e-05 | 6.7895 | 0.0599 | [8 empty] | > 95% (spike) | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0339 | 87205 | 806.93 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1224 | 51778 | 1823 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM1 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0246 | 90080 | 662 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0343 | 87660 | 813.64 |  |  |  |  | 99.9 |  | 27 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1232 | 52084 | 1833 |  |  |  |  | 99.9 |  | 21 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Naneos |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 | Partector | Y | Y | http://www.naneos.ch/ | Electrical | sensor system |  | PM2.5 | #/L vs µm2/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0249 | 90539 | 668 |  |  |  |  | 99.9 |  | 28 | [38 empty] | 7000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/naneos-partector---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8267 | 4.123 | 0.6655 |  |  |  |  | 99 |  | 16 | [14 empty] | -70 - 91.2 over 0-180 ug/m3 | [23 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8152 | 4.1645 | 0.6718 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8274 | 4.202 | 0.7498 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9299 | 3.5897 | 0.6942 |  |  |  |  | 99 |  | 16 | [14 empty] | -70 - 91.2 over 0-180 ug/m3 | [23 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9245 | 3.4702 | 0.7101 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9323 | 3.6368 | 0.7838 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8971 | 3.0295 | 0.578 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8896 | 2.988 | 0.5956 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.889 | 3.1992 | 0.6448 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.92 | -35 | 1.6 |  |  |  | 22.3 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.92 | -35 | 1.8 |  |  |  | 22.3 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.96 | 2.9 | 2 |  |  |  | 52.7 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Airviz |  | Carnegie Mellon | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.96 | 5.3 | 2.3 |  |  |  | 52.7 | [41 empty] | 150 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.29 | 6.7 | 0.74 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.21 | 5.4 | 1.06 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | min | [17 empty] | GRIMM | Sensor | linear | 0.07 | 5.9 | 0.4 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.28 | 7.7 | 0.55 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.46 | 4.3 | 1.9 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Airviz |  | Carnegie Mellon | Zikova et al. 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.15 | 5.7 | 0.66 | [27 empty] | 9 | [17 empty] | 150 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Alphasense |  | CitiSense | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NO2-A1 | Y | Y |  | electrochemical | OEM | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.9772 | [16 empty] | 34.2ppb @ > 200ppb SO2 |  |  | 4.6 | SD | [21 empty] | YES |  |  |  |  | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NORMAL | JRC-FK
Alphasense |  | CitiSense | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | NO2-A1 | Y | Y |  | electrochemical | OEM | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.9722 | [16 empty] | 34.2ppb @ > 200ppb SO2 |  |  | 23.3 | SD | [21 empty] | YES |  |  |  |  | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HUMID | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9183 | 2.8934 | 0.5874 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9119 | 2.836 | 0.6061 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9148 | 3.0078 | 0.6587 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9509 | 2.6128 | 0.603 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6284 | 2.426 | 0.6284 |  |  |  |  | 99 |  | 16 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6792 | 2.6785 | 0.6792 |  |  |  |  | 99 |  | 14.2 | [38 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---field-evaluation.pdf | PM evaluation Kit | JRC-FK
Shinyei |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---lab-evaluation.pdf | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.93 | -18.46 | 0.84 | [7 empty] | 20 | 40 | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/shinyei-pm-eval-kit---lab-evaluation.pdf | PM evaluation Kit | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | 0.805 | -5.472 | 9.2455 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.644 | -4.1935 | 9.341 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 7.5 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | 0.883 | 3.89 | 1.0306 |  |  |  |  | log | PM=a1*ln(Rs)+a0 | [8 empty] | Sensor | GRIMM | [8 empty] | 100 | 100 | 10 | [33 empty] | YES |  | YES |  |  | 310 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n2---field-evaluation.pdf?sfvrsn=0 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppb | Y | black box | LAB |  |  | 60 | 1 | min | 0.9909 | -11.51 | 1.3291 |  |  |  |  | linear | [9 empty] | Sensor | REF | linear | [48 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/aeroqual-s500-o3---lab-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8386 | 0.0018 | 0.98029 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/aeroqual-s500-o3---lab-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8554 | 0.0013 | 1.0061 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8606 | 6e-05 | 1.0064 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8571 | 0.0014 | 0.991 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8743 | 0.001 | 1.0157 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
City University of Hong Kong |  | HKEPD/14-02771 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppm/mV | Y | black box | LAB |  |  |  | 1 | min | 0.99 | -15.14 | 1.08 |  |  |  |  | linear | CO=a0+a1*Rs | [8 empty] | Sensor | REF | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [25 empty] | NO |  | NO |  |  |  | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
City University of Hong Kong |  | HKEPD/14-02772 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - NO2-B4 | NO2 | ppb/mV | Y | black box | LAB |  |  |  | 1 | min | 0.99 | -0.99 | 0.0026 |  |  |  |  | linear | NO2=a0+a1*Rs | [8 empty] | Sensor | REF | [44 empty] | YES | [5 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
City University of Hong Kong |  | HKEPD/14-02773 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - CO-B4 | CO | ppm/mV | Y | black box | FIELD | urban |  |  | 1 | min |  | 0.099 | 0.0025 |  |  |  |  | linear | CO=a0+a1*Rs | [8 empty] | REF | Sensor | linear | 0.97 | [12 empty] | 0.085 |  |  | -0.053 | -0.034 | [25 empty] | NO |  | NO |  |  |  | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
City University of Hong Kong |  | HKEPD/14-02774 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N |  | electrochemical | sensor system | AlphaSense - NO2-B4 | NO2 | ppb/mV | Y | black box | FIELD | urban |  |  | 1 | min |  | -3.96 | 10.97 | -0.17 | 0.33 |  |  | MLR | NO2=(Rs-a1*RH-a0)/(a2*RH+a3) |  |  | RH | [5 empty] | REF | Sensor | linear | 0.9 | [42 empty] | YES | [5 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
2B Technologies |  | HKEPD/14-02774 | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | N | N | https://twobtech.com/products.html | UV | sensor system | 2B Tech. (POM) | O3 | ppb/mV | Y | open source | LAB |  |  | 20 | 1 | min | 0.99 | -0.99 | 0.0026 |  |  |  |  | linear | O3=a0+a1*Rs | [8 empty] | Sensor | REF | [50 empty] | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.471 | 455.38 | 0.6977 | -1.0835 |  |  |  | MLR | NO2=a0+a1*Rs1+a2*Rs2 | [8 empty] | REF | Sensor | linear | 0.288 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.612 | 715.45 | 0.8394 | -1.4811 |  | 0.5326 |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a4*RH |  |  | RH | [5 empty] | REF | Sensor | linear | 0.293 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.556 | 827.92 | 0.8688 | -1.5498 | -1.6344 |  |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T |  | T | [6 empty] | REF | Sensor | linear | 0.821 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.617 | 790.88 | 0.8707 | -1.5645 | -0.5051 | 0.4513 |  | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T+a4*RH |  | T | RH | [5 empty] | REF | Sensor | linear | 0.82 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
KNMI |  | https://making-sense.eu/ | Mijling et al. 2017 | 2017 | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf | Urban AirQ | N | N |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  |  | 1 | hour | 0.788 | 274.85 | 0.3186 | -0.4805 | -0.5447 | -0.4744 | -0.5349 | MLR | NO2=a0+a1*Rs1+a2*Rs2+a3*T+a4*RH+a5*O3 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.887 |  |  | 5.31 | [45 empty] | USD | https://www.atmos-meas-tech.net/11/1297/2018/amt-11-1297-2018.pdf |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf | Intel Berkeley Badge | Y | N |  | electrochemical | sensor system | City Technology - NO2 3E50 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.9862 | 1622.4 | 0.0857 |  |  |  |  | linear | y = a0 + a1 x | [8 empty] | REF | Sensor | linear | [49 empty] | USD | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf | Intel Berkeley Badge | Y | N |  | electrochemical | sensor system | City Technology - NO2 3E50 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.9823 | 1628 | 0.0948 |  |  |  |  | linear | y = a0 + a1 x | [8 empty] | REF | Sensor | linear | [49 empty] | USD | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf |  | JRC-FK
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf | Intel Berkeley Badge | Y | N |  | MOs | sensor system | SGX Sensortech - MICS-4514 | NO2 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.8913 | 690.16 | 488.47 |  |  |  |  | log | Y = a0 + a1 ln(x) | [8 empty] | REF | Sensor | logaritmic | [49 empty] | USD | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.001 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.068 | 1727 | -247.53 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.002 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.51 | -1977.94 | 280.68 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | ANN | All parameters: CO sensors, T | [8 empty] | REF | Sensor | linear | 0.8 |  | 0.62 | 0.05 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.768 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.078 | -78.28 | 1.097 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.562 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.062 | -108.18 | 16.01 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | O3 | V vs ppb | Y | open source | LAB |  | 108 |  | 1 | hour | 0.9986 | 0.031 | 0.00485 | [10 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.76 | -0.011 | 7e-05 |  |  | 1.3 | -0.022 | 0.6 | SD | [13 empty] | 1.6 |  | 1.2 | < 0.142 | [9 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - O3-3E1F | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.845 | [6 empty] | linear | Rs=a1*O3+a0 | [8 empty] | REF | Sensor | linear | 0.667 | -26.31 | 1.84 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - O3-3E1F | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.845 | [6 empty] | linear | Rs=a1*O3+a0 | [8 empty] | REF | Sensor | linear | 0.667 | -26.31 | 1.84 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.91 | 46 | 0.89 |  |  |  |  | RF | [9 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - O3-3E1F | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.852 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.584 | -30.15 | 2.1 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - O3-3E1F | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.945 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.824 | -12.03 | 1.49 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -15.91 | 0.1593 |  | 0.7 |  |  | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.9606 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.84 | 10 | 0.82 | 7.96 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -15.93 | 0.1606 | -0.00175 | 0.7 | 0.000121 | 1.87e-07 | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8836 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.85 | 12.6 | 0.84 | 10.33 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -0.09 | 0.1547 |  | 0.63 |  |  | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8464 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.87 | 7.4 | 0.85 | 6.14 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3783/2017/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B42F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min |  | -0.13 | 0.1548 | -0.000204 | 0.63 | 0.000134 | 1.81e-07 | MLR | Rs=a0+a1*NO2+a2*NO+a3*RH+a4*CO+a5*CO2+a5*O3 | O3 | 0.8464 | RH | NO |  | CO | CO2 |  | REF | Sensor | linear | 0.8 | 11.1 | 0.83 | 9.23 | [44 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/10/3783/2017/ | coefficient units = (mv ppb-1) | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.9 |  |  | 4.12 | [44 empty] | 3538 | EUR | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf |  | JRC-FK
decentLab |  | OpenSense II | Mueller et al. 2017 | 2017 | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-OX-B431 | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.97 |  |  | 2.05 | [44 empty] | 3538 | EUR | https://static1.squarespace.com/static/5612610de4b0a8aca36ee66a/t/5645cbb4e4b0b229a0675e3f/1447414708764/Decentlab+Air+Quality.pdf |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO)*T | [8 empty] | REF | Sensor | linear | 0.9 |  |  | 4.6 | 3.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
Figaro |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | MOs | OEM | Figaro TGS-5042 | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | [7 empty] | ANN | CO = f(CO-TGS5042, Absolute Humidity, MICS-4514-NO2, Temperature, CO2 712) | [8 empty] | REF | Sensor | linear | 0.351 | 0.04 | 1.38 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | ANN using raw data | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.1 | 2.3 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO2 = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO2)*T | [8 empty] | REF | Sensor | linear | 0.6 |  |  | 5.6 | 4.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.95 |  |  | 3.3 | 2.1 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 7.1 | 4.2 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO)*T | [8 empty] | REF | Sensor | linear | 0.8 |  |  | 8.3 | 5.8 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO-B4 | NO | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 7.1 | 4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-B4 | NO | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | ANN | All parameters: NO, NO2, O3, CO sensors, T, RH, Press | [8 empty] | REF | Sensor | linear | 0.97 |  | 0.97 | 4.4 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | NO2 = a0+a1*Rs(NO)+a2*Rs(NO2)+a3*T+a4*RH+a5*RS(NO2)*T | [8 empty] | REF | Sensor | linear | 0.76 |  |  | 5.5 | 4.4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | [7 empty] | ANN | NO = f(MICS-2710, Relative Humidity, Temperature, NO-3E100) | [8 empty] | REF | Sensor | linear | 0.18 | 0.7 | 0.57 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | City Technology - sensoric - NO2-3E50 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | [7 empty] | ANN | NO2 = f(NO2 3E50, MICS-4514-NO2, O3 3E1F, MICS-2710, MICS-4514-CO, NO2B4 107, absolute humidity) | [8 empty] | REF | Sensor | linear | 0.596 | 5.81 | 0.64 | [45 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | ANN with raw data | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | electrochemical | sensor system | A3CO-Citytech | CO | mg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.8 |  | 1.01 | 0.066 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | electrochemical | sensor system | KWJ-Engineering-SNL-NO2 | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.71 |  | 1.01 | 4.76 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
AirVeraCity |  | OpenSense II | Marjovi et al. 2017 | 2017 | https://dl.acm.org/citation.cfm?id=3108012 | AirVeraCity | Y | Y | https://airveracity.com/ | MOs | sensor system | SGX-SensorTech-MiCS_OZ 47 | O3 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.75 |  | 1.01 | 5.08 | [44 empty] | 10000 | EUR | https://dl.acm.org/citation.cfm?id=3108012 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD1 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.81 | 11.542 | 0.727 | 0.097 | -0.437 | 1.5416 |  | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | MLR (R), Support Vector Machines, e1017 (R), ANN - Neuralnet (R), RandomForest (R), R caret PACKAGES | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.74 |  |  | 4.6 | 3.5 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.79 |  |  | 4.1 | 3.2 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | SVR | [9 empty] | REF | Sensor | linear | 0.83 |  |  | 4.6 | 3.6 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD2 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.93 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
decentLab |  | OpenSense II | Bigi et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/3717/2018/ | AirCube | Y | Y | https://www.decentlab.com/air-quality-monitoring/ | electrochemical | sensor system | Alphasense-NO2-B43F | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | min | [7 empty] | RF | [9 empty] | REF | Sensor | linear | 0.85 |  |  | 4.4 | 3.4 | [43 empty] | 3538 | EUR | https://www.atmos-meas-tech.net/11/3717/2018/ |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD1 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.9 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.91 |  |  | 4.88 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD1 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.81 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.92 |  |  | 4.09 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD3 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.85 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD1 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.88 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.36 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD2 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.86 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 8.67 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD2 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.89 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.95 |  |  | 8.26 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD4 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.9 | [6 empty] | MLR | NO2=a0+a1*Rs+a2*NO+a3*O3+a4*T | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | [48 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD2 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.62 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.98 |  |  | 6.34 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD3 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.85 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.91 |  |  | 4.58 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD3 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.79 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.92 |  |  | 4.16 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8792 | -0.0003 | 1.0145 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8686 | 0.0013 | 0.9939 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8922 | 0.0008 | 1.0202 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 | S-500 | Y | Y | https://www.aeroqual.com/product/series-500-portable-ozone-gas-monitor | MOs | sensor system | S-500 OZU | O3 | ppm | Y | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.8962 | -0.0004 | 1.0185 |  |  |  |  | 100 |  | 0.03 | [38 empty] | 500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-s500-ozl---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | N | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.826 | 12.306 | 0.0379 | -0.977 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/unitec-sens-it-co---lab-evaluation.pdf?sfvrsn=8 | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | LAB |  |  | 60 | 1 | min | [17 empty] | Sensor | FRM | linear | 0.9877 | -3.6185 | 3.0707 | [40 empty] | YES |  | YES |  |  | 2200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/unitec-sens-it-co---lab-evaluation.pdf?sfvrsn=8 |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.426 | 0.1956 | -6e-05 | 0.343 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.15 | [14 empty] | 36.5 -50 over 0-25 ppm | [23 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.3431 | 0.2233 | 0.2085 | 0.1547 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.6 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.3299 | 0.2289 | 0.2031 | 0.1667 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 0.6 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.6017 | 2.2345 | 0.0012 | 0.1682 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 21 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.5678 | 2.2187 | 0.0012 | 0.1309 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 25 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.6207 | 1.2418 | 0.001 | 0.1689 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 26 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.628 | 2.2107 | 0.0011 | 0.1727 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 21 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.5977 | 2.2026 | 0.0012 | 0.1327 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 25 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.6475 | 1.217 | 0.0009 | 0.172 |  |  |  | quad | NO2=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 99 |  | 26 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.716 | 9.6771 | 0.0195 | -0.004 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | 0.7676 | 9.6273 | 0.0186 | -0.0701 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.8301 | 12.322 | 0.0381 | -0.9875 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7221 | 9.2623 | 0.0194 | 0.0049 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7726 | 9.2604 | 0.0186 | -0.00621 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.7229 | 3.0643 | 0.0182 | 0.1326 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 35 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.6343 | 3.3406 | -0.0124 | 1.6537 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 30 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
UNITEC |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf | SENS-IT | Y | Y | http://www.unitec-srl.com/site/products/sens-it/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.6798 | -1.8537 | -0.0069 | 1.3516 |  |  |  | quad | O3=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | REF | [8 empty] | 99 |  | 31 | [38 empty] | 2200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/unitec-sens-it---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.4855 | -0.0743 | -0.2945 | 1.3282 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.8303 | 0.171 | -0.8806 | 0.0481 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | 0.7751 | -0.033 | 0.1717 | 0.871 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.487 | -0.0758 | -0.3303 | 1.3512 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.8437 | 0.1693 | 0.8764 | 0.0577 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | 0.7939 | 0.0034 | 0.6935 | 0.6935 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.4531 | -0.0571 | -0.5204 | 1.4293 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.814 | 3107 | 0.2113 | 0.2448 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.808 | 18.9 | 0.94 | 26.6 | [16 empty] | 140 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.949 | 2385 | 0.4851 | -1.574 | 0.1029 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.916 | 0.6 | 0.96 | 6.1 | [16 empty] | 22 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.881 | 6142 | -0.1479 | -0.1559 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.839 | -2.4 | 1.04 | 7.7 | [16 empty] | 35 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.914 | 6135 | -0.2636 | -0.2597 | 0.1249 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.84 | -5.9 | 1.1 | 9.8 | [16 empty] | 37 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.903 | 3124 | 0.233 | 0.1435 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.872 | 12 | 0.94 | 21 | [16 empty] | 120 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.95 | 2382 | 0.4505 | -1.039 | 0.08234 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.889 | 2.5 | 0.96 | 7 | [16 empty] | 25 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.872 | 6146 | -0.1409 | -0.05854 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * RH | [8 empty] | REF | Sensor | linear | 0.814 | 1.1 | 1.02 | 8 | [16 empty] | 40 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.902 | 6140 | -0.2222 | -0.2196 | -0.03732 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor,f(T)) + a3 * T | [8 empty] | REF | Sensor | linear | 0.878 | -7.6 | 1.02 | 7.6 | [16 empty] | 22 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.923 | 2379 | 0.4731 | -0.6332 | 0.07075 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.898 | 0.2 | 0.96 | 6.9 | [16 empty] | 27 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.88 | 6145 | -0.149 | -0.1037 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.833 | 0.2 | 0.96 | 7.3 | [16 empty] | 30 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.91 | 6137 | -0.2335 | -0.2364 | -0.1559 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.85 | -4.3 | 1.03 | 8.7 | [16 empty] | 24 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.871 | 3115 | 0.2221 | 0.2002 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.858 | 13.5 | 0.95 | 22.3 | [16 empty] | 120 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.948 | 2386 | 0.4462 | -1.024 | 0.08168 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.93 | 0.3 | 0.98 | 5.4 | [16 empty] | 24 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.871 | 6141 | -0.1544 | -0.2648 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.816 | -1.3 | 0.97 | 7.8 | [16 empty] | 30 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.908 | 6140 | -0.1993 | -0.2273 | -0.04756 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.862 | -6.4 | 1.16 | 9.5 | [16 empty] | 50 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.899 | 3133 | 0.2388 | 0.05537 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.88 | 5.1 | 0.96 | 20.7 | [16 empty] | 120 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.938 | 2384 | 0.468 | -1.228 | 0.08926 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.922 | 0.2 | 0.98 | 5.9 | [16 empty] | 24 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.833 | 6139 | -0.1405 | -0.5694 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.799 | -3.5 | 1.09 | 9.2 | [16 empty] | 55 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | min | 0.917 | 6138 | -0.2393 | -0.23 | -0.01232 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.894 | -6 | 1.09 | 8.4 | [16 empty] | 30 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (1 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.814 | 3107 | 0.2113 | 0.2448 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.864 | 19.5 | 0.94 | 19.7 | [16 empty] | 98 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.949 | 2385 | 0.4851 | -1.574 | 0.1029 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.9332 | 0.3 | 0.98 | 4.9 | [16 empty] | 19 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.881 | 6142 | -0.1479 | -0.1559 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.9215 | -4.3 | 1.09 | 4.9 | [16 empty] | 18 | U at 80 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_01 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.914 | 6135 | -0.2636 | -0.2597 | 0.1249 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.875 | -7.2 | 1.13 | 8.7 | [16 empty] | 30 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_01 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.903 | 3124 | 0.233 | 0.1435 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.91 | 14.6 | 0.93 | 16.5 | [16 empty] | 102 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.95 | 2382 | 0.4505 | -1.039 | 0.08234 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.889 | 2.4 | 0.97 | 6.3 | [16 empty] | 15 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.872 | 6146 | -0.1409 | -0.0585 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * RH | [8 empty] | REF | Sensor | linear | 0.908 | -0.4 | 1.06 | 5.1 | [16 empty] | 18 | U at 80 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_02 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.902 | 6140 | -0.2222 | -0.2196 | -0.03732 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor, f(T)) + a3 * T | [8 empty] | REF | Sensor | linear | 0.91 | -7.6 | 1.01 | 6.4 | [16 empty] | 17 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_02 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.923 | 2379 | 0.4731 | -0.6332 | 0.07075 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.911 | -0.6 | 0.97 | 5.6 | [16 empty] | 18 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.88 | 6145 | -0.149 | -0.1037 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.91 | -1.3 | 1.01 | 4.9 | [16 empty] | 9 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_03 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.91 | 6137 | -0.2335 | -0.2364 | -0.1559 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.891 | -4.9 | 1.04 | 7.3 | [16 empty] | 22 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_03 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.872 | 3115 | 0.2221 | 0.2002 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.89 | 17 | 0.94 | 17.1 | [16 empty] | 75 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.95 | 2386 | 0.4462 | -1.024 | 0.08168 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.948 | -0.1 | 1 | 4.2 | [16 empty] | 17 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.871 | 6141 | -0.1544 | -0.2649 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.898 | -3 | 1.02 | 5.2 | [16 empty] | 11.5 | U at 80 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_04 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.908 | 6140 | -0.1993 | -0.2273 | -0.04756 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.901 | -7.3 | 1.18 | 7.8 | [16 empty] | 35 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_04 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.899 | 3133 | 0.2388 | 0.05537 |  |  |  | MLR | Rs = a0 + a1 * CO + a2 * RH | [8 empty] | REF | Sensor | linear | 0.911 | 6.8 | 0.96 | 15.6 | [16 empty] | 95 | U at 500 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.938 | 2384 | 0.468 | -1.228 | 0.08926 |  |  | MLR | Rs = a0 + a1 * NO + a2 * T + a3 * T² | [8 empty] | REF | Sensor | linear | 0.941 | 0 | 0.98 | 4.6 | [16 empty] | 13 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.833 | 6139 | -0.1405 | -0.5694 |  |  |  | MLR | Rs = a0 + a1 * NO2 + a2 * T | [8 empty] | REF | Sensor | linear | 0.889 | -5.5 | 1.15 | 6.2 | [16 empty] | 25 | U at 80 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio |  | Proof of Concept AirSensEUR | Karagulian et al. 2019 | 2019 | AirSensEUR-NL_05 | AirSensEUR (v.2) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | 0.917 | 6138 | -0.2393 | -0.23 | -0.01232 |  |  | MLR | Rs = a0 + a1 * O3 + a2 * NO2(sensor) + a3 * T | [8 empty] | REF | Sensor | linear | 0.928 | -6.6 | 1.11 | 6.3 | [16 empty] | 24 | U at 100 | [26 empty] | 1600 | EUR | AirSensEUR-NL_05 | summer field test (60 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | linear | NO = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.79 |  | 0.96 | 12.5 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO-A4 | NO | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | MLR | NO = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.93 |  | 0.99 | 6.7 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD3 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.88 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.96 |  |  | 3.4 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | linear | NO2 = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.74 |  | 0.99 | 9 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | MLR | NO2 = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.81 |  | 0.97 | 7.1 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD4 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.95 | [6 empty] | RF | [9 empty] | REF | Sensor | linear | 0.97 |  |  | 7.44 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | linear | O3 = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.79 |  | 0.93 | 10.5 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | MLR | O3 = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.8 |  | 0.96 | 10.3 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD4 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.95 | [6 empty] | SVM | [9 empty] | REF | Sensor | linear | 0.97 |  |  | 5.68 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | linear | CO = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.81 |  | 0.79 | 0.04 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense CO-A4 | CO | µg/m3 | Y | open source | FIELD | traffic |  | 60 | 1 | hour | [7 empty] | MLR | CO = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.81 |  | 0.78 | 0.04 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
AQmesh |  |  | Cordero et al. 2018 | 2018 | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 | AQMesh v.4.0 | Y | Y | POD4 | electrochemical | sensor system | AlphaSense | NO2 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | 0.86 | [6 empty] | ANN | [9 empty] | REF | Sensor | linear | 0.93 |  |  | 7.09 | [44 empty] | 10000 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0925400518307159 |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_b | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | linear | NO2 = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.93 |  | 0.85 | 7.2 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (15 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_b | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | MLR | NO2 = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.93 |  | 0.86 | 7 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (15 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | ANN | All parameters: NO, NO2, O3, CO sensors, T, RH, Press | [8 empty] | REF | Sensor | linear | 0.9 |  | 0.89 | 4.7 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_b | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | linear | CO = a0 + a1 * Rs | [8 empty] | REF | Sensor | linear | 0.9 |  | 0.63 | 100 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (15 min. data) | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_b | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | MLR | CO = a0 + a1 * Rs + a2 * RH + a3*T | [8 empty] | REF | Sensor | linear | 0.9 |  | 0.6 | 110 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (15 min. data) | JRC-FK
Samyoung |  |  | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | DSM501A | Y | Y |  | nephelometer | OEM | DSM501A | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2.5 | hour | 0.8914 | 3.93 | 59.7 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 19 | [5 empty] | 11.4 | 4.28 | [8 empty] | YES | 100 | YES | 250 | YES | 100 | YES | 15 |  | 15 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
Samyoung |  |  | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | DSM501A | Y | Y |  | nephelometer | OEM | DSM501A | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 24 | hour | 0.5 | [6 empty] | linear | [9 empty] | Sensor | DustTrak | [49 empty] | 15 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.8205 | 0.1187 | 0.5324 | 0.4258 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.56 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Acrobotic |  |  | AQ-SPEC | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf | Smart Citizen Kit | Y | Y |  | MOs | sensor system | SGX-Sensortech - MICS-4514 | CO | ppm | N | black box | FIELD | rural |  | 60 | 8 | hour | 0.771 | 0.0014 | 0.3249 | 0.7138 |  |  |  | quad | CO=a1*Rs2+a2*Rs+a0 | [8 empty] | Sensor | FRM | [8 empty] | 96 |  | 0.58 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/smart-citizen-kit---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.003 | 0.4134 | -0.1364 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0013 | 0.4121 | -0.0723 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | CO | ppm | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0025 | 0.4129 | -0.088 |  |  |  |  | 100 | [40 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0277 | 20.345 | -0.5731 |  |  |  |  | 100 |  | 0 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0207 | 14.535 | 0.1351 |  |  |  |  | 100 |  | 41 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | NO2 | ppb | N | black box | FIELD | rural | 30 | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0095 | 14.868 | 0.0674 |  |  |  |  | 100 |  | 79 | [38 empty] | 240 | USD | http://db-airmontech.jrc.ec.europa.eu/download/air-quality-egg-v2_co-no2---field-evaluation.pdf |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0 | 0 | 0.1253 |  |  |  |  | 100 |  | 90 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1442 | 30.209 | -0.081 |  |  |  |  | 100 |  | 160 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 | Egg v.2 | Y | Y | https://airqualityegg.com/egg | electrochemical | sensor system | Spec Sensors | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1737 | 31.343 | -0.0786 |  |  |  |  | 100 |  | 190 | [38 empty] | 240 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_ozone---field-evaluation.pdf?sfvrsn=2 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.823 | 3.5456 | 0.9781 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8365 | 3.2115 | 1.0069 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8351 | 3.5039 | 0.983 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8488 | 3.1674 | 1.0119 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9282 | 3.44 | 0.9876 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
TSI |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.64 | 2.64 | 1.14 | 0.41 |  |  |  | 87 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.608 | -0.04 | 1.13 |  |  |  |  | 87 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.81 | -1.38 | 1.19 |  |  |  |  | 74 | [40 empty] | 1500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TSI |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.45 | 107 | 3.1 |  |  |  | 13.4 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
TSI |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -7.3 | 7.7 |  |  |  | 2.6 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
TSI |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.42 | 122 | 2.8 |  |  |  | 4 | [41 empty] | 1500 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban | 60 |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  | 0.81 | [11 empty] | 0.012 | -0.0002 | 1 | [19 empty] | 0.94 | [12 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban | 60 |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 |  | 0.81 | [11 empty] | 0.012 | -0.002 | 1 | [19 empty] | 0.94 | [12 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.99 | [12 empty] | 0.2 | 0.24 | [34 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | good sensor-to-sensor reproducibility | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - NO-AF | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.8 | [12 empty] | 0.0045 |  | 0.001 | [33 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Mead et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S1352231012011284 | SNAQ | Y | Y |  | electrochemical | sensor system | AlphaSense - NO-AF | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.95 | [12 empty] | 0.0045 |  | 0.001 | [33 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231012011284 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 16.82 | 1.01 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - CO-AF | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 2.53 | 0.96 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.74 | 2.35 | 1 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.83 | 7.76 | 0.97 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.92 | 2.53 | 0.96 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.78 | -0.94 | 0.94 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.71 | 3.23 | 0.93 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 | -0.27 | 0.9 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO-A1 | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.84 | 2.48 | 0.86 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.88 | 5.83 | 0.81 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
University of Cambridge |  | SNAQ Heathrow | Popoola et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016308317 | SNAQ | N | Y |  | electrochemical | sensor system | AlphaSense - NO2-A1 | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.85 | 3.35 | 0.82 | [42 empty] | YES |  |  |  | USD | https://www.sciencedirect.com/science/article/pii/S1352231016308317 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM2.5 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9469 | 3.1393 | 1.0138 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3126 | 15.781 | 1.1348 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0682 | 21.925 | 0.8345 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3461 | 15.784 | 1.135 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3713 | 15.106 | 1.2003 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0659 | 22.221 | 0.7287 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4418 | 15.703 | 1.1281 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 | Egg v.2 (PM) | Y | Y | https://airqualityegg.com/egg | nephelometer | sensor system | Shinyei-PPD60 | PM10 | µg/m3 | N | black box | FIELD | rural | 30 |  | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4776 | 15.048 | 1.1921 |  |  |  |  | 100 | [37 empty] | YES |  |  | 280 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg-v2_pm---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=9 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.65 | 9.318 | 0.138 |  |  |  |  | 96.3 | 100 |  | 20 | 40 | [12 empty] | -1086-(-408) over 10-50 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=9 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.66 | 9.021 | 0.18 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.66 | 9.7253 | 0.1038 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.685 | 9.2526 | 0.1405 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.695 | 8.94 | 0.183 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6918 | 9.659 | 0.1053 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.737 | 8.819 | 0.15 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.744 | 8.5842 | 0.1956 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7415 | 9.225 | 0.113 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7365 | 82050 | 7.907 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.74 | 77809 | 9.43 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7458 | 89100 | 6.55 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.758 | 80754 | 8 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.763 | 76425 | 9.55 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.769 | 87820 | 6.6426 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.89 | 61779 | 9.379 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9972 | [16 empty] | 19 ppb NO2 at > 200 ppb SO2 |  |  | 1.2 | SD | [12 empty] | ppb | [14 empty] | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9919 | [19 empty] | 6.4 | SD | [12 empty] | ppb | [14 empty] | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9937 | [19 empty] | 7.4 | SD | [12 empty] | ppb | [14 empty] | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
AGT |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | AGT ATS-35 NO2 | N | N |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9917 | [19 empty] | 7.5 | SD | [12 empty] | ppb | [14 empty] | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.893 | 57023 | 11.165 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9 | 70633 | 7.74 |  |  |  |  | 80 | [40 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.659 | 10.359 | 0.175 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.668 | 9.707 | 0.2317 |  |  |  |  | 80 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6622 | 10.908 | 0.1307 |  |  |  |  | 80 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.767 | 9.499 | 0.1985 |  |  |  |  | 80 |  | 40.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.77 | 8.988 | 0.2599 |  |  |  |  | 99 |  | 33 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7684 | 10.098 | 0.148 |  |  |  |  | 99 |  | 29 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airbeam---field-evaluation.pdf?sfvrsn=4 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.87 | -8.52 | 0.16 |  |  |  |  | 100 |  |  | 20 | 40 | [36 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 |  | JRC-FK
HabitatMap |  | AirCasting | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 | AirBeam | Y | Y | http://www.takingspace.org/aircasting/airbeam/ | OPC | sensor system | Plantower-PMS7003 | PM2.5 | #/L | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.77 | -3000000 | 5.66 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -1086 | -408 | % |  |  | 0 | [18 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/airbeam---laboratory-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0628 | 596.21 | -0.1267 |  |  |  |  | 5.5 |  | 1050 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8397 | 170.6 | 0.6855 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0751 | 779.71 | 0.2104 |  |  |  |  | 10.2 |  | 650 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8779 | 120.46 | 0.7236 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8726 | 170.33 | 0.6814 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS011 | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 2 | min | 0.8634 | 6.3 | 2 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | REF | Sensor |  |  |  |  | 2.6 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS011 | PM10 | µg/m3 | N | black box | LAB |  |  |  | 2 | min | 0.7679 | 5 | 0.7 |  |  |  |  | linear | PM10=a0+a1*Rs | [8 empty] | REF | Sensor |  |  |  |  | 3.3 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS011 | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 24 | hour | [17 empty] | REF | Sensor | linear | 0.957 | 0.777 | 0.959 | 4.056 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
CNR (IT) |  |  | Cavaliere et at. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2843 | AIRQino | N | Y | https://pan.baidu.com/s/1skCk8AL | OPC | sensor system | NovaFitness - SDS011 | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 24 | hour | [17 empty] | REF | Sensor | linear | 0.909 | 1.288 | 0.972 | 6.084 | [44 empty] | 1000 | EUR | https://www.mdpi.com/1424-8220/18/9/2843 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | OPC | sensor system | University of Hertfordshire, CAIR | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.15 |  |  |  | 125.1 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.06 |  |  |  | 20.17 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | MOs | sensor system | Sensortech SGX e2v MiCS-2710 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.76 |  |  |  | 0.09 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | electrochemical | sensor system | Alphasense - O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.13 |  |  |  | 22.12 | [43 empty] | 280 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8988 | 124.51 | 0.7126 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8926 | 132.59 | 0.688 |  |  |  |  | 77.3 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | CO | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.9181 | 94.594 | 0.7484 |  |  |  |  | 89.2 |  | 500 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.1379 | 28.252 | -0.3172 |  |  |  |  | 96.7 |  | 24 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.1601 | 34.701 | -0.2835 |  |  |  |  | 99.5 |  | 55 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0004 | 6.1318 | -0.0094 |  |  |  |  | 0.4 |  | 17 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0586 | 27.758 | -0.1567 |  |  |  |  | 98.5 |  | 55 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.0611 | 10.819 | 0.1644 |  |  |  |  | 86 | 97 | 38 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.2363 | 5.7843 | 0.1953 |  |  |  |  | 86 | 97 | 60 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.0061 | 14.377 | 0.0396 |  |  |  |  | 86 | 97 | 60 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
Spec Sensors |  | Developer Kit | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf | Spec | Y | Y |  | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1453 | 8.4929 | 0.1627 |  |  |  |  | 86 | 97 | 58 | [38 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/spec-sensors---field-evaluation249fa8efc2b66f27bf6fff00004a91a9.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8137 | 2.589 | 0.6731 |  |  |  |  | 99 |  | 16 | [14 empty] | -108- (-7.2) over 0-150 ug/m3 | [23 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8304 | 0.9655 | 0.6959 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.81 | 0.9805 | 0.7248 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8299 | 2.4787 | 0.6797 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8466 | 0.8409 | 0.7033 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8297 | 0.7892 | 0.736 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8716 | 2.629 | 0.6768 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8779 | 0.9703 | 0.7064 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8632 | 0.9374 | 0.7403 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7647 | 2.1327 | 0.6852 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7734 | 0.5203 | 0.7062 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.757 | 0.4767 | 0.7394 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8737 | 2.1322 | 0.6705 |  |  |  |  | 99 |  | 16 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8853 | 0.3528 | 0.7035 |  |  |  |  | 99 |  | 18 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8876 | 0.2103 | 0.7443 |  |  |  |  | 99 |  | 17 | [38 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
TSI |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf | AirAssure | Y | Y |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -5.37 | 0.5 |  |  |  |  | 97 | [40 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/tsi-airassure---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.755 | 5.054 | 0.2685 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7645 | 4.901 | 0.2905 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7858 | 4.5891 | 0.3309 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7678 | 5.0321 | 0.2689 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7755 | 4.917 | 0.289 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7973 | 4.604 | 0.3303 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6279 | 6.69 | 0.221 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.666 | 6.384 | 0.244 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7091 | 6.023 | 0.282 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.653 | 3.891 | 0.345 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.621 | 4.692 | 0.355 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 4.2059 | 0.4089 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.659 | 4.3277 | 0.326 |  |  |  |  | 80 |  | 38 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6318 | 5.2236 | 0.33 |  |  |  |  | 91 |  | 30 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Hanvon |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf | Hanvon N1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6659 | 4.8 | 0.3827 |  |  |  |  | 91 |  | 33 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/hanvon---field-evaluation.pdf |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=6 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.573 | 6.087 | 0.43 |  |  |  |  | 96.3 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=7 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.546 | 5.71 | 0.512 |  |  |  |  | 96.3 |  | 18 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=7 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=8 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.547 | 4.514 | 0.42 |  |  |  |  | 96.3 |  | 22 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=10 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.657 | 6.511 | 0.414 |  |  |  |  | 96.3 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=11 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.583 | 6.58 | 0.462 |  |  |  |  | 96.3 |  | 18 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=11 |  | JRC-FK
Foobot |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=12 | Foobot | Y | Y | https://foobot.io/features/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.568 | 6.07 | 0.359 |  |  |  |  | 96.3 |  | 22 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/foobot---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5711 | 5.883 | 0.4958 |  |  |  |  | 75 |  | 15 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5724 | 6.274 | 0.5414 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5862 | 6.2765 | 0.566 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0285 | 40.083 | 0.332 |  |  |  |  | 75 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0509 | 38.737 | 0.411 |  |  |  |  | 99.9 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.045 | 39.159 | 0.4191 |  |  |  |  | 99.9 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6636 | 6.3369 | 0.4788 |  |  |  |  | 75 |  | 15 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7664 | 6.0143 | 0.5643 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7498 | 6.2295 | 0.5746 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.01 | 43.548 | 0.1482 |  |  |  |  | 75 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0252 | 42.181 | 0.2293 |  |  |  |  | 99.9 |  | 19.5 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Origins |  |  | AQ-SPEC | 2016 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 | Laser Egg | Y | Y | Shinyei | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0258 | 42.083 | 0.2525 |  |  |  |  | 99.9 |  | 19 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/laser-egg---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9384 | -0.9237 | 0.9738 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO2-B42F (688150) | NO2 | ppb | N | black box | LAB |  | <25 |  | 15 | sec | 0.98 | -1.02 | 1.22 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | 10 | 63 | [8 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense CO-B4 (688150) | CO | ppb | N | black box | LAB |  | <15 |  | 15 | sec | 0.98 | 0.07 | 0.86 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.44 | 10 | 63 | [5 empty] | -0.034 |  |  | 25 | SD (ppb) | [25 empty] | 6 months - significant change | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO2-B42F (864150) | NO2 | ppb | N | black box | LAB |  | <25 |  | 15 | sec | 0.922 | 3.85 | 1.21 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | 10 | 63 | [8 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO2-B42F (688150) | NO2 | ppb | N | black box | FIELD | urban | <25 |  | 15 | sec | 0.423 | 3.8 | 0.38 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 30.27 | [5 empty] | 0.46 | [10 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense CO-B4 (688150) | CO | ppb | N | black box | FIELD | urban | <15 |  | 15 | sec | 0.336 | 166 | 0.88 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 170.99 | [5 empty] | 0.44 | [7 empty] | -0.034 |  |  | 25 | SD (ppb) | [25 empty] | 6 months - significant change | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO2-B42F (864150) | NO2 | ppb | N | black box | FIELD | urban | <25 |  | 15 | sec | 0.09 | 16 | 0.2 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.46 | [10 empty] | 2.8 | SD | [25 empty] | significant change during 6 months | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense OX-B421 (688150) | O3 | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | -1.27 | 1.16 |  |  |  |  | linear | [5 empty] | NO2 |  |  |  | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense OX-B421 (864150) | O3 | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | 3.25 | 0.99 |  |  |  |  | linear | [5 empty] | NO2 |  |  |  | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense OX-B421 (688150) | O3 | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.81 | 7.2 | 0.26 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear |  |  |  | 22.2 | [5 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense OX-B421 (864150) | O3 | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.32 | 9 | 0.11 |  |  |  |  | linear | [9 empty] | REF | Sensor | linear | [9 empty] | 0.32 | [38 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO-B4(688150) | NO | ppb | N | black box | LAB |  |  |  | 15 | sec | 0.9801 | 0.07 | 0.86 |  |  |  |  | linear | [15 empty] | 16.35 | [44 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Castell et all. 2017 | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO-B4(688150) | NO | ppb | N | black box | FIELD | urban |  |  | 15 | sec | 0.9801 | -1.13 | 0.97 |  |  |  |  | linear | [60 empty] | 10000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S0160412016309989-main.pdf |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  |  |  | 1.46 | [43 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.86 |  |  |  | 0.05 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [30 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7 |  |  |  | 2.4 | [43 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AQmesh |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AQMesh v.4.0 | Y | Y |  | electrochemical | sensor system | AlphaSense NO-A4 | NO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [20 empty] | 0.8 |  |  |  | 1.51 |  |  |  | 37 | [39 empty] | 10000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9369 | 0.3975 | 0.8723 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9467 | -0.0896 | 0.9071 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9025 | 2.3264 | 0.7228 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9087 | 3.0881 | 0.676 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9193 | 2.7581 | 0.6902 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3933 | 12.85 | 0.6651 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3969 | 13.378 | 0.6206 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.4124 | 12.932 | 0.6467 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9467 | -1.089 | 0.9825 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9405 | 0.3316 | 0.8757 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.952 | -0.1427 | 0.912 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9128 | 2.2696 | 0.7271 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9168 | 3.0625 | 0.6783 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9279 | 2.7253 | 0.6927 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4282 | 12.876 | 0.6646 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.432 | 13.416 | 0.6195 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4483 | 12.981 | 0.6448 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9884 | -0.8786 | 0.9695 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9834 | 0.4733 | 0.8623 |  |  |  |  | 99.9 |  | 10 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9884 | 0.054 | 0.894 |  |  |  |  | 99.9 |  | 10.5 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9676 | 2.4996 | 0.7134 |  |  |  |  | 99.9 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9682 | 3.3068 | 0.6635 |  |  |  |  | 99.9 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9739 | 3.0201 | 0.6747 |  |  |  |  | 99.9 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4779 | 13.357 | 0.6485 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4799 | 13.932 | 0.6017 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4962 | 13.542 | 0.6244 |  |  |  |  | 99.9 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7754 | 3.2227 | 0.7099 |  |  |  |  | 87 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7773 | 3.995 | 0.6637 |  |  |  |  | 87 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7944 | 3.6326 | 0.68 |  |  |  |  | 87 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9067 | 2.8604 | 0.7409 |  |  |  |  | 87 |  | 13 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9168 | 3.6259 | 0.6961 |  |  |  |  | 87 |  | 14 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9281 | 3.3161 | 0.7081 |  |  |  |  | 87 |  | 15 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3219 | 25.189 | 0.6845 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3264 | 25.704 | 0.6407 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3424 | 25.208 | 0.6708 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4578 | 24.94 | 0.7197 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4688 | 25.448 | 0.6773 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.4864 | 25.01 | 0.7026 |  |  |  |  | 87 |  | 17 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purpleair---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.95 | -20.2 | 0.76 |  |  |  |  | 74.5 | 100 |  | 20 | 40 | [12 empty] | -4.1-77.2 over 15-300 ug/m3 | [23 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.997 | -2.41 | 0.35 |  |  |  |  | 100 | 100 |  | 20 | 40 | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2016 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf | PA-I | Y | N | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | LAB |  |  | 20 | 1 | min | [17 empty] | Sensor | GRIMM | linear | 0.96 | -1.96 | 0.57 |  |  |  |  | 66 | 100 |  | 20 | 40 | [36 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-i---lab-evaluation.pdf |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.878 | -0.96 | 0.906 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.85 | -0.56 | 0.683 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.421 | 5.768 | 0.365 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.806 | 3.075 | 0.599 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.782 | 3.49 | 0.489 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.591 | 5.854 | 0.375 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.156 | 26.86 | 0.506 |  |  |  |  | 66 | 86 | 30 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.198 | 23.97 | 0.46 |  |  |  |  | 66 | 86 | 39 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.227 | 22.8 | 0.457 |  |  |  |  | 66 | 86 | 38 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | -0.8463 | 0.896 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | -0.6169 | 0.685 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6896 | 1.6093 | 0.592 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.821 | 2.948 | 0.599 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.805 | 3.479 | 0.488 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.773 | 3.12 | 0.483 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.94 | 0.8 | 0.94 |  |  |  |  | linear | Rs=a0+a1*NO2 | [8 empty] | REF | Sensor | linear | 0.69 | 3.26 | 0.81 | 4.56 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.96 | 9.19 | 0.96 |  |  |  |  | linear | Rs=a0+a1*CO | [5 empty] | CO |  |  | REF | Sensor | linear | 0.88 | 3.98 | 0.94 | 32.9 | [9 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - NO-B4 | NO | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.94 | 0.46 | 0.94 |  |  |  |  | linear | Rs=a0+a1*NO | [8 empty] | REF | Sensor | linear | 0.84 | 1.29 | 0.94 | 4.52 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
Alphasense |  |  | Cross et al. 2017 | 2017 | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf | ARISense | Y | Y |  | electrochemical | sensor system | Alphasense - OX-B421 | O3 | ppb | N | black box | FIELD | urban |  |  | 5 | min | 0.65 | 8.34 | 0.65 |  |  |  |  | linear | Rs=a0+a1*O3 | [8 empty] | REF | Sensor | linear | 0.39 | 13.1 | 0.47 | 9.71 | [45 empty] | USD | https://www.atmos-meas-tech.net/10/3575/2017/amt-10-3575-2017.pdf |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2017 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | NO2 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.88 | 3.251 | 0.9407 | 9.29 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2018 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | O3 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.91 | 0.2077 | 1.0039 | 6.52 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
Bettair Cities |  |  | Bettair 2017 | 2018 | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ | Bettair | Y | Y | https://bettaircities.com/ | electrochemical | sensor system |  | NO | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.95 | 0.7292 | 0.9974 | 10.05 | [44 empty] | 1000 | EUR | https://bettaircities.com/how-bettair-is-solving-the-cities-air-pollution-crisis/ |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.94 | -1.07 | 0.89 |  |  |  |  | 66 | 86 | 18 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.93 | -1.049 | 0.694 |  |  |  |  | 66 | 86 | 20.5 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.961 | -0.9156 | 0.719 |  |  |  |  | 66 | 86 | 20 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.88 | 2.12 | 0.613 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.867 | 2.57 | 0.507 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.933 | 1.843 | 0.526 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.812 | 4.36 | 0.6206 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.802 | 4.904 | 0.506 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.731 | 5.303 | 0.501 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.88 | 4.06 | 0.629 |  |  |  |  | 66 | 86 | 22 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.86 | 4.351 | 0.523 |  |  |  |  | 66 | 86 | 24 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.89 | 3.75 | 0.55 |  |  |  |  | 66 | 86 | 23 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
Envea |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | LAB |  | 2305 |  | 1 | hour | 0.9996 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.007 | -0.001 | -0.01 |  | 0.093 | -0.057 | 1.7 | SD | [13 empty] | 4.8 |  | 1.6 | < -0.065 | [9 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | semi-rural |  |  | 1 | hour | 0.465 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.004 | 438.99 | -60.94 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | semi-rural |  |  | 1 | hour | 0.24 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.036 | 229.86 | -29.14 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | semi-rural |  |  | 1 | hour | 0.745 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.021 | 16.17 | -97.84 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | semi-rural |  |  | 1 | hour | 0.585 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.004 | 32.5 | -205.92 | [45 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.28 | 1.91 | 1.16 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.01 | 13.6 | 0.04 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.13 | 10.92 | 1.28 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip NO2-F | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.28 | 24.23 | 1.57 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol Membrapor O3/M-5 | O3 | ppb | Y | black box | LAB |  | 150 |  | 1 | hour | 0.9973 | [12 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.84 | 0.006 | 0.07 |  |  | 0.24 |  | 3.3 | SD | [13 empty] | 9.2 |  | 1.4 | < 0.011 | [9 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9968 | [28 empty] | 7 |  |  |  | 3.8 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9978 | [28 empty] | 0 |  |  |  | 2.9 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9975 | [28 empty] | 8.6 |  |  |  | 4.2 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol - Membrapor NO2/M-20 | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9896 | [28 empty] | 8.3 |  |  |  | 9.3 | ppb | [13 empty] | 600 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol Membrapor O3/M-5 | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.79 | 4.79 | 0.93 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Duvall at al. 2016 | 2016 | https://www.mdpi.com/1424-8220/16/10/1698 | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol Membrapor O3/M-5 | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.72 | 3.35 | 1.1 | [45 empty] | 600 | EUR | https://www.mdpi.com/1424-8220/16/10/1698 |  | JRC-FK
Envea |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol Membrapor O3/M-5 | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7569 | -10 | 0.65 |  |  |  |  | 63 | [40 empty] | 600 | EUR | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Envea |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | CairClip O3/NO2 | Y | Y | http://cairpol.com/en/downloads/technical-resources/ | electrochemical | sensor system | CairPol Membrapor O3/M-5 | O3 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7056 | -15 | 0.67 |  |  |  |  | 82 | [40 empty] | 600 | EUR | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Envea |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CairClip PM2.5 | Y | Y | https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance | OPC | sensor system | CairPol | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0117 | 0.1311 | -0.0053 | [7 empty] | > 19.8 | > 95% (spike) | [36 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Envea |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CairClip PM2.5 | Y | Y | https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance | OPC | sensor system | CairPol | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0637 | 2.6107 | -0.2289 | [7 empty] | > 19.8 | > 95% (spike) | [36 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
ENEA |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | AirSensorBox | N | Y |  | nephelometer | sensor system | shinyei PPD20V | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.33 |  |  |  | 21.15 | [43 empty] | 30 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | OPC | sensor system | University of Hertfordshire, CAIR | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.13 |  |  |  | 80.91 |  |  |  | 93 | [40 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | OPC | sensor system | University of Hertfordshire, CAIR | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.07 |  |  |  | 17.15 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.84 |  |  |  | 5.61 |  |  |  | 93 | [40 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 |  |  |  | 0.18 |  |  |  | 92 |  |  |  |  | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
University of Cambridge |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | CAM | N | Y |  | electrochemical | sensor system | AlphaSense NO-A4 | NO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [20 empty] | 0.34 |  |  |  | 12.02 |  |  |  | 93 | [40 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
AirBase |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CanarIT | Y | N |  | OPC | sensor system |  | PM | µg/m3 | N | black box | FIELD | rural |  | 20 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.0394 | 6.5448 | -0.1023 | [45 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
AirBase |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | CanarIT | Y | N |  | OPC | sensor system |  | PM | µg/m3 | N | black box | FIELD | rural |  | 20 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0044 | 6.4979 | -0.1013 | [45 empty] | 1500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.258 | 33.69 | 0.639 |  |  |  |  | 66 | 86 | 30 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.258 | 32.435 | 0.549 |  |  |  |  | 66 | 86 | 39 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
AS-LUNG |  | Air Quality Station | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 | Air Quality Station | N | Y | http://www.plantower.com/en/content/?107.html | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.279 | 31.546 | 0.604 |  |  |  |  | 66 | 86 | 38 | [38 empty] | 2000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-air-quality-station---field-evaluation18ce99efc2b66f27bf6fff00004a91a9.pdf?sfvrsn=14 |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.3116 | -1.1229 | 0.5626 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0006 | 12.106 | -0.0029 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.2708 | -0.5655 | 0.497 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | CO 3E300 | Y | Y |  | electrochemical | OEM | CityTechnology - Sensoric - CO 3E300 | CO | V/ppm | Y | open source | LAB |  |  |  | 1 | hour | 1 | 1.66 | 0.01959 |  |  |  |  | linear | Rs=a0+a1*CO |  |  |  | NO |  |  |  |  | REF | Sensor | [14 empty] | -0.007 ± 111 | 0.50 ± 0.080 |  | -0.74 ± 0.54 | -3.1 ± 5.3 | -25.6 ± 1.4 | 0.26 ± 0.82 | [7 empty] | 0.056 | ppm.min |  |  | 0-500 | ppb | [10 empty] | YES |  | YES |  |  | 100 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.96 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.09 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.98 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.03 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.97 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.03 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | CO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | CO | ppm | Y | open source | LAB |  |  |  | 1 | hour | 0.98 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [14 empty] | 0.01 | 0.01 |  | 0.01 |  |  |  | 0.02 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 2 | hour | 0.88 | [6 empty] | quad | Rs=(1+0.25*RH2)/(1-RH) |  |  | RH | [5 empty] | BAM | Sensor | linear | 0.75 | -3.22 | 1.22 | [38 empty] | YES |  |  | YES |  |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | RH < 60% | JRC-FK
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [11 empty] | RH | [5 empty] | PARTISOL | Sensor | linear | 0.93 | 1.85 | 0.92 | [38 empty] | YES |  |  | YES |  |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 |  | JRC-FK
Thermo Electron Corp |  |  | Chakrabarti et al. 2004 | 2004 | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 | DataRAM pDR-1200 | Y | N |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [11 empty] | RH | [5 empty] | BAM | Sensor | linear | 0.71 | 1.06 | 0.61 | [38 empty] | YES |  |  | YES |  |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S1352231004001955 |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.1206 | 28085 | 211.16 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.82 | -1370 | 1.35 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.89 | -3799 | 1.81 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.72 | -1300 | 1.3 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.85 | -2299 | 1.61 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Testo |  |  | Viana et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231015304921 | DiscMini | Y | Y |  | OPC | OEM |  | PM | # | Y | open source | FIELD | urban |  | 600 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.75 | -432 | 1.16 | [45 empty] | 11000 | USD | https://www.sciencedirect.com/science/article/pii/S1352231015304921 |  | JRC-FK
Sharp |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | DN7C3CA006 | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -56 | 1.3 |  |  | -6.3 | [16 empty] | 4.2 | [16 empty] | YES | 120 | [7 empty] | 10 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Dylos |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 | PM2.5-0.5 | #/ft3 | N | open source | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.533 | 86531 | 21562 | [8 empty] | > 95% (increase) | [31 empty] | YES |  | NO |  |  | 300 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Dylos |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 | PM2.5-0.5 | #/ft3 | N | open source | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.5483 | 51574 | 21368 | [8 empty] | > 95% (increase) | [31 empty] | YES |  | NO |  |  | 300 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.7396 | -152 | 64 | 0.15 |  |  |  | 82 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.6084 | -1182 | 428 |  |  |  |  | 72 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5329 | -941 | 431 |  |  |  |  | 77 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0008 | 147465 | 2.0367 |  |  |  |  | 95 |  | 980 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0532 | 52685 | 144.9 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.3288 | -2.0194 | 0.6009 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0008 | 12.127 | -0.0036 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2844 | -1.3926 | 0.5326 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM10-2.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.16 | 5.5 | 1.3 | 0.1 |  |  |  | 82 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM10-2.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.1089 | 73 | 5.7 |  |  |  |  | 72 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM10-2.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0729 | 84 | 4.9 |  |  |  |  | 77 | [40 empty] | 300 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.88 | 67 | 0.95 |  |  |  | 13.4 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | 0.96 | 30.1 | 0.0056 | 1.99 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | [49 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM10 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.67 | 77.65 | 0.7 |  |  |  | 13.4 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.94 | 0.91 | 0.1 |  |  |  | 46.1 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | 0.82 | 0.42 | -0.000154 | 0.15 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | linear | [48 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1100 PRO | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1100 PRO | PM10 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.89 | 2.1 | 0.094 |  |  |  | 46.1 | [41 empty] | 300 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.87 | 73 | 0.89 |  |  |  | 13.4 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | 0.93 | 38.86 | 0.0052 | 1.86 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | [49 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.7 | 82.21 | 0.68 |  |  |  | 13.4 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.93 | 0.99 | 0.11 |  |  |  | 46.1 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | 0.78 | 0.42 | -0.000184 | 0.17 |  |  |  | quad | Rs=a0+a1*PM2.5²+a2*PM2.5 | [8 empty] | REF | Sensor | linear | [48 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 |  | JRC-FK
Dylos |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10 | #/cm3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.85 | 2.27 | 0.1 |  |  |  | 46.1 | [41 empty] | 475 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Dylos |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10-2.5 | #/cm3 | Y | open source | LAB |  |  |  | 1 | min | 0.98 | 98 | 8.6 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | SMPS | Sensor | linear | 0.98 | 10 | 1.2 |  |  | -18 |  |  |  | 3.1 | [13 empty] | 8 | [15 empty] | YES | 350 | YES | [6 empty] | 475 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/m3 | Y | open source | LAB |  |  |  | 1 | min | 0.98 | -43.6 | 0.018 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | DustTrak | linear | [30 empty] | 1 | [17 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Smoke and PLS | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/m3 | Y | open source | FIELD | traffic |  |  | 24 | hour | [17 empty] | Sensor | DustTrak | linear | 0.99 | 0.008 | 0.51 | [45 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0021850216301835?via%3Dihub |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM3 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.98 |  |  | 0.8 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.99 |  |  | 0.6 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.78 |  |  | 4 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5 | #/cm3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.58 |  |  | 3.5 | [44 empty] | 475 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Dylos |  |  | Steinle et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.9 | 1.29 | 1.11e-05 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | TEOM | [10 empty] | 6.5 | [29 empty] | NO | [8 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub |  | JRC-FK
Dylos |  |  | Steinle et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | hour | 0.7 | 4.75 | 2.8e-05 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | TEOM-FDMS | [10 empty] | 6.5 | [29 empty] | NO | [8 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0048969714017057?via%3Dihub |  | JRC-FK
Dylos |  |  | Han et al,.2016 | 2016 | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | min | 0.778 | -2.8197 | 0.0071 |  |  |  |  | linear | PM2.5=a0+a1*Rs | [8 empty] | Sensor | GRIMM | [10 empty] | 11.3 | [31 empty] | YES | 300 | YES | 110 |  |  |  | 475 | USD | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true |  | JRC-FK
Dylos |  |  | Han et al,.2016 | 2016 | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10-2.5 | #/ft3 | Y | open source | FIELD | urban |  |  | 1 | min | 0.481 | -1.0689 | 0.0544 |  |  |  |  | linear | PM10-2.5=a0+a1*Rs | [8 empty] | Sensor | GRIMM | [10 empty] | 4.8 | [31 empty] | YES | 400 | YES | 150 |  |  |  | 475 | USD | https://www.tandfonline.com/doi/pdf/10.1080/10962247.2016.1241195?needAccess=true |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD |  |  |  | 1 | min | 0.94 | 0.4922 | 0.7895 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10-2.5 | #/ft3 | Y | open source | FIELD |  |  |  | 1 | min | 0.77 | 0.01422 | 0.9114 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | #/ft3 | Y | open source | FIELD |  |  |  | 1 | min | 0.439 | -4.053 | 3.417 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10-2.5 | #/ft3 | Y | open source | FIELD |  |  |  | 1 | min | 0.00499 | 0.6443 | -0.001876 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | TSI OPS | [49 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour |  | 0.55 | 6.926 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | GRIMM | linear | 0.84 | [47 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
Dylos |  |  | Jovasevic_Stojanovic et al 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115300294 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM10 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour |  | 1.603 | 3.357 |  |  |  |  | linear | PM=a0+a1*Rs | [8 empty] | Sensor | GRIMM | linear | 0.73 | [47 empty] | 475 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115300294 |  | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_b | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | ANN | All parameters: NO2 sensors, T | [8 empty] | REF | Sensor | linear | 0.94 |  | 0.78 | 7.5 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (15 min. data) | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract | BAIRS | N | N | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | mg/m3 | N | open source | LAB |  |  |  |  | hour | [7 empty] | linear | [9 empty] | Sensor | DustTrack |  | 0.99 | 0.008 | 0.51 | [26 empty] | < 1 | [18 empty] | 475 | USD | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract |  | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract | BAIRS | N | N | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | mg/m3 | N | open source | LAB |  |  |  |  | hour | [7 empty] | linear | [9 empty] | Sensor | DustTrack |  | 0.81 | 0.027 | 0.57 | [26 empty] | < 1 | [18 empty] | 475 | USD | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract |  | JRC-FK
Dylos |  |  | Northcross et al. 2013 | 2013 | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract | BAIRS | N | N | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5-0.5 | mg/m3 | N | open source | LAB |  |  |  |  | hour | [7 empty] | linear | [9 empty] | Sensor | DustTrack |  | 0.98 | 0.0052 | 1.049 | [26 empty] | < 1 | [18 empty] | 475 | USD | https://pubs.rsc.org/en/content/articlelanding/2013/em/c2em30568b#!divAbstract |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.1264 | 19465 | 226.47 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0009 | 147152 | 2.2379 |  |  |  |  | 95 |  | 980 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | counts/L | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.053 | 48225 | 152.68 |  |  |  |  | 95 |  | 500 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2486 | -0.8308 | 0.6312 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | Shinyei -shinyei-PPD42 | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.33 |  |  |  | 24.57 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | Shinyei -shinyei-PPD42 | PM10 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.36 |  |  |  | 23.61 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | shinyei-PPD42 | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.23 |  |  |  | 11.26 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | nephelometer | sensor system | shinyei-PPD42 | PM2.5 | µg/m3 | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.27 |  |  |  | 11.11 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | electrochemical | sensor system | CityTechnology-NO2 3E50 | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 |  |  |  | 4.95 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
ECN |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ECN_Box | N | Y |  | electrochemical | sensor system | Alphasense-O3-B4 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.14 |  |  |  | 21.5 | [43 empty] | 274 | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Sensaris |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Eco PM | N | N |  | OPC | sensor system |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 30 | sec | [17 empty] | GRIMM | Sensor | linear | 0.315 | 0.0037 | 0.034 | [46 empty] | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 7e-06 | 13.908 | -0.0004 |  |  |  |  | 95 |  | 38 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
Airviz |  | Carnegie Mellon | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf | Speck | Y | Y |  | nephelometer | sensor system | DSM501A | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.1482 | 1.678 | 0.472 |  |  |  |  | 95 |  | 21 | [38 empty] | 150 | USD | http://db-airmontech.jrc.ec.europa.eu/download/speck-sensor---field-evaluation.pdf |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.868 | -0.872 | 0.877 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.866 | -0.789 | 0.888 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8661 | -1.2827 | 0.842 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.78 | 3.257 | 0.602 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7758 | 3.385 | 0.621 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7803 | 2.88 | 0.589 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.113 | 27.65 | 0.546 |  |  |  |  | 87 |  | 26 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.131 | 26.616 | 0.585 |  |  |  |  | 87 |  | 26 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.137 | 23.503 | 0.533 |  |  |  |  | 87 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.888 | -0.8897 | 0.876 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Perkin Elmer |  |  | US-EPA_2015 | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF | ELM | N | N | https://scistarter.com/tools/776 | MOs | sensor system |  | NO2 | ppb | N | black box | FIELD | rural |  |  | 1 | hour | [17 empty] | CAPS | Sensor | linear | 0.004 | 18.552 | 0.808 | [8 empty] | > 92% (spike) | [36 empty] | 5200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SPECKANDELMEVAL-EPAREPORT-NOVEMBER2-2015.PDF |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8857 | -0.8077 | 0.887 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8838 | -1.2893 | 0.8404 |  |  |  |  | 82 |  | 15 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.798 | 3.27 | 0.6015 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | sensor system | OPC-N2 | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | SidePak | linear | 0.497 |  |  | 2.94 | [45 empty] | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 |  | JRC-FK
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppm | N | black box | FIELD | indoor |  |  | 5 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.99 |  |  | 0.01792 | [41 empty] | YES |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 |  | JRC-FK
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | black box | FIELD | indoor |  |  | 5 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.93 |  |  | 3.204 | [41 empty] | YES |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 |  | JRC-FK
Alphasense |  |  | Gillooly et al. 2019 | 2019 | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 | EMMA | N | Y |  | electrochemical | sensor system | Alphasense NO-B4 | NO | ppb | N | black box | FIELD | indoor |  |  | 1 | min | [10 empty] | T | [6 empty] | Sensor | REF | linear | 0.97 |  |  | 3.221 | [41 empty] | YES |  |  |  | USD | https://linkinghub.elsevier.com/retrieve/pii/S0269749118319584 |  | JRC-FK
FIS |  |  | Spinelle et al. 2016 | 2016 | https://www.aidic.it/cet/16/54/054.pdf | FIS SP-61 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | LAB |  |  | 89 | 1 | hour | 0.9991 | 10.8 | 2.86 | 0.00235 | -0.000109 |  |  | MLR | Rs=a0+a1*O3+a2*O3²+a3*O33 | [8 empty] | REF | Sensor | linear | [48 empty] | 50 | USD | https://www.aidic.it/cet/16/54/054.pdf |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7961 | 3.44 | 0.617 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.792 | 2.9267 | 0.588 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.936 | -1.289 | 0.8795 |  |  |  |  | 82 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.936 | -1.306 | 0.8947 |  |  |  |  | 82 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.932 | -1.8971 | 0.8538 |  |  |  |  | 82 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.905 | 2.206 | 0.6217 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Sharp |  |  | Olivares and Edwards, 2015 | 2015 | https://www.atmos-meas-tech-discuss.net/amt-2015-175/ | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [7 empty] | linear | PM2.5=0.6*Rs+19; Rs=-4.4*T - 49 |  | T | [6 empty] | TEOM-FDMS | Sensor | linear | 0.72 |  |  | 14.1 | [41 empty] | NO |  |  | 10 | USD | https://www.atmos-meas-tech-discuss.net/amt-2015-175/ |  | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.98 | 3.24 | 0.45 |  |  |  | 46.1 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.94 | 3.4 | 0.3 |  |  |  | 96.9 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.98 | 1.56 | 0.41 |  |  |  | 118.2 | [41 empty] | 10 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | Arizona Test Dust | JRC-FK
Sharp |  |  | Sousan et al 2016 | 2016 | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.99 | -55 | 1.3 |  |  | -5.8 | [16 empty] | 3.1 | [16 empty] | YES | 300 | [7 empty] | 10 | USD | http://www.tandfonline.com/doi/full/10.1080/02786826.2016.1162901 | Arizona Road Dust | JRC-FK
Sharp |  |  | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | mg/m3/Volts | Y | open source | LAB |  |  |  | 1 | sec | 0.9844 |  | -4e-05 | 0.2227 |  |  |  | quad | Rs=a1*PM2.52+a2Rs | [8 empty] | Sensor | DustTrak | [49 empty] | 10 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Sharp |  |  | Alvarado et al. 2015 | 2015 | https://www.mdpi.com/1424-8220/15/8/19667 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM10 | mg/m3/Volts | Y | open source | LAB |  |  |  | 1 | sec | 0.6461 |  | -0.0002 | 0.2908 |  |  |  | quad | Rs=a1*PM102+a2Rs | [8 empty] | Sensor | DustTrak | [49 empty] | 10 | USD | https://www.mdpi.com/1424-8220/15/8/19667 |  | JRC-FK
Sharp |  |  | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | GP2Y1010AU0F | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 2.5 | hour | 0.9838 | 91.1 | 196 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 25 | [5 empty] | 26.1 | 26.9 | [8 empty] | YES | 1000 | YES | 0.75 | YES | 40 | YES | 15 |  | 10 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
Aristotle University |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ISAG | N | N |  | MOs | sensor system | SGX-Sensortech - MiCS-2710 | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.02 |  |  |  | 16.17 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Aristotle University |  |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | ISAG | N | N |  | MOs | sensor system | SGX-Sensortech - MICS-2610 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.12 |  |  |  | 360.12 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.87 | -3.1 | 0.95 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.89 | 2.1 | 0.94 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9 | 0.29 | 0.95 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.96 | 1.7 | 0.94 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.93 | 0.93 | 0.96 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | NO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.93 | 3.2 | 0.92 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.43 | 1.26 | 0.55 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR A10 V2 | Y | Y | https://www.kunak.es/en/ | electrochemical | sensor system | AlphaSense | CO | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.27 | 243 | 0.41 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR P10 V2 | Y | Y | https://www.kunak.es/en/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.54 | -0.33 | 0.66 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
kunak |  |  | Kunak field report 2017 | 2017 | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ | KUNAKAIR P10 V2 | Y | Y | https://www.kunak.es/en/ | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.82 | -4 | 0.76 | [45 empty] | 5000 | EUR | https://www.kunak.es/en/products/ambient-monitoring/particle-counter/ |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.901 | 2.3672 | 0.6391 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9021 | 1.8564 | 0.3074 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.83 | 4.091 | 0.638 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.827 | 4.233 | 0.657 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.825 | 3.731 | 0.623 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.912 | 3.808 | 0.645 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9122 | 3.9416 | 0.6652 |  |  |  |  | 85 |  | 21 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 15 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9106 | 3.426 | 0.6315 |  |  |  |  | 85 |  | 22 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0967 | 40365 | 0.593 |  |  |  |  | 87 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0952 | 40.97 | 0.605 |  |  |  |  | 87 |  | 25 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
AS-LUNG |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 | Portable AS-LUNG | N | Y | Portable | OPC | sensor system | Plantower - PMS3003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 15 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0933 | 40.477 | 0.557 |  |  |  |  | 87 |  | 28 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/as-lung-portable---field-evaluation.pdf?sfvrsn=8 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8489 | 9.12 | 0.731 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | LAB |  |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a1*exp(a2*CO)*(T-298)+a3*exp(a4*CO)+a5*exp(a6*CO)*RH) |  | T | RH | [11 empty] | 3.56 | [32 empty] | 0.49 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a0+a1*CO+a2*T+a3*RH+a4*time) |  | T | RH | [11 empty] | 0.44 | [32 empty] | 0.49 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | LAB |  |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a1*exp(a2*NO2)*(T-298)+a3*exp(a4*NO2)+a5*exp(a6*NO2)*RH) |  | T | RH | [11 empty] | 7.1 | [32 empty] | 0.93 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | min | [7 empty] | MLR | Rs=Rs0*(a0+a1*NO2+a2*T+a3*RH+a4*time) |  | T | RH | [11 empty] | 8.4 | [32 empty] | 0.93 |  | -0.01 | [6 empty] | YES |  |  |  | USD | https://www.atmos-meas-tech.net/7/3325/2014/ | Rs is the resistance of the MOS sensor | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | N | N |  | MOs | sensor system | MiCS-5525 - SGX-Sensortech | CO | ppm | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.16 | [8 empty] | 307.1 | [12 empty] | 0.31 | [26 empty] | USD | https://www.atmos-meas-tech.net/7/3325/2014/ |  | JRC-FK
University of Colorado Boulder |  |  | Piedrahita et al. 2014 | 2014 | https://www.atmos-meas-tech.net/7/3325/2014/ | M-POD | Y | N |  | MOs | sensor system | MiCS-2710 - SGX-Sensortech | NO2 | ppb | N | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | Sensor | FEM | linear | 0.06 | [8 empty] | 11.3 | [12 empty] | 2.46 | [26 empty] | USD | https://www.atmos-meas-tech.net/7/3325/2014/ |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | GRIMM | Sensor | linear | 0.2307 | -0.1285 | 0.1301 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.6368 | -0.0562 | 0.057 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Met One |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | Met One - 831 | Y | Y |  | OPC | sensor system |  | PM10 | counts | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7729 | -0.008 | 0.0493 | [8 empty] | > 90% (spike) | [36 empty] | 2050 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.884 | 9.274 | 0.8265 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.812 | 10.873 | 0.77 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8533 | 7.9017 | 0.793 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8707 | 8.2385 | 0.889 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8529 | 8.901 | 0.8785 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7 | 3.5 | 0.844 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.592 | 4.98 | 0.8647 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.744 | 4.228 | 0.985 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | Sun et al. 2016 | 2016 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ | MAS | Y | Y |  | OPC | sensor system | Met One-ES642 | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | min |  |  | 0.75 | 0.25 |  |  |  | MLR | PM2.5=Rs*a1*(1+a2(RH2)/(1-RH)) |  |  | RH | [5 empty] | REF | Sensor | linear | 0.92 | [47 empty] | 5500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC4801587/ |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8 | 6.44 | 0.825 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.802 | 7.0186 | 0.919 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.781 | 7.369 | 0.909 |  |  |  |  | 20 | 23 | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.759 | 3.31 | 0.37 |  |  |  |  | 20 | 23 | 15 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.658 | 4.72 | 0.858 |  |  |  |  | 20 | 23 | 16 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Moji China |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 | AirNut | Y | Y |  | OPC | sensor system | Plantower | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8 | 4.31 | 0.97 | [6 empty] | 12 | [38 empty] | 150 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-nut---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6919 | 6.8046 | 0.601 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6933 | 6.7944 | 0.5973 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7224 | 5.9376 | 0.5485 |  |  |  |  | 98 |  | 21 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.798 | 8.2503 | 0.519 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7989 | 8.1951 | 0.5186 |  |  |  |  | 98 |  | 17 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.808 | 7.485 | 0.4769 |  |  |  |  | 98 |  | 21 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.249 | 23.604 | 0.9069 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2494 | 23.863 | 0.8989 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4076 | 16.185 | 0.888 |  |  |  |  | 98 |  | 31 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.1636 | 30.622 | 0.613 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.573 | 3.778 | 1.1198 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7951 | -0.0068 | 1.2038 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7186 | -8.491 | 1.1094 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8713 | -7.5874 | 1.496 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8758 | -2.1549 | 1.3772 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.8036 | -1.4041 | 1.1638 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.7294 | -9.3819 | 1.4734 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.6246 | -2.1884 | 1.619 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
RTI |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | MicroPEM | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.572 | -2.8132 | 1.2934 | [8 empty] | > 94% (spike) | [36 empty] | 2000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.206 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.131 | 8.13 | -58.83 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.2 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.126 | 8.48 | -56.45 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.745 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a0 | O3 | T | [6 empty] | REF | Sensor | linear | 0.021 | 3176.44 | -477.45 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MiCS-2710 | Y | N |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.744 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a0 | O3 | T | [6 empty] | REF | Sensor | linear | 0.063 | -207890.74 | 31407.64 | [45 empty] | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensortech |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | MiCS-2710 | Y | N |  | MOs | OEM | SGX Sensortech - MiCS-2710 | NO2 | ppb | N | open source | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9851 | [19 empty] | 5 | SD | [12 empty] | ppb | [8 empty] | YES |  |  |  |  | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HOT | JRC-FK
SGX-Sensortech |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | MiCS-2710 | Y | N |  | MOs | OEM | SGX Sensortech - MiCS-2710 | NO2 | ppb | N | open source | LAB |  | 60 |  | 1 | hour | [20 empty] | 0.9815 | [19 empty] | 2.6 | SD | [12 empty] | ppb | [8 empty] | YES |  |  |  |  | 7 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | HUMID | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.78 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.066 | -1.48 | 9.64 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.78 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.067 | -1.76 | 11.03 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.76 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.035 | -2.92 | 17.19 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensortech |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | MICS-4514 | Y | Y |  | MOs | OEM |  | CO | ppm | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.77 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.047 | -2.59 | 15.19 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.168 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.016 | 56.37 | -378.51 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.269 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.203 | 6.18 | -40 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.525 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*NO+a4*T+a0 | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | 0.01 | -3970.53 | 577.84 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
SGX-Sensotech |  |  | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | MICS-4514 | Y | Y |  | MOs | OEM | e2V, MiCS | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.786 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*NO+a4*T+a0 | O3 | T |  | NO |  |  |  |  | REF | Sensor | linear | 0.016 | 5692.9 | -852.21 | [45 empty] | 20 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Envira | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | Y | Y | https://enviraiot.com/nanoenvi-and-its-applications-for-air-quality-control/ | electrochemical | sensor system | Alphasense - NO2-B43F | NO2 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.57 |  |  |  | 14.93 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Envira | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | Y | Y | https://enviraiot.com/nanoenvi-and-its-applications-for-air-quality-control/ | MOs | sensor system | SGX-Sensotech-MICS-OZ-47 | O3 | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.77 |  |  |  | 7.66 | [44 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 |  | JRC-FK
Envira | NILU |  | Borrego et al. 2016 | 2016 | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | NanoEnvi | N | Y | https://enviraiot.com/nanoenvi-and-its-applications-for-air-quality-control/ | electrochemical | sensor system | Alphasense - CO-B4 | CO | ppb | N | black box | FIELD | traffic |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.53 |  |  |  | 0.1 | [8 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.sciencedirect.com/science/article/pii/S1352231016307610 | cross sensitivity come from Lewis et al. 2016 | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.0057 | [6 empty] | linear | Rs = a1*NO+a0 | [8 empty] | REF | Sensor | linear | 0.018 | -315.8 | 307 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.62 | 4.5 | 0.72 |  |  |  |  | RF | [9 empty] | REF | Sensor | [50 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 2.3e-05 | [6 empty] | linear | Rs = a1*NO+a0 | [8 empty] | REF | Sensor | linear | 0.001 | -21860 | 23106 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.0018 | [6 empty] | MLR | Rs = a1*NO+a2*T+a3*RH+a0 | [8 empty] | REF | Sensor | linear | 0.009 | 972 | -1056.6 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | ppb | Y | open source | FIELD | semi-rural | < 20 |  | 1 | hour | 0.075 | [6 empty] | MLR | Rs = a1*NO+a2*T+a3*RH+a0 | [8 empty] | REF | Sensor | linear | 0.02 | 118.62 | 108.9 | [45 empty] | 120 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | NO-3E100 | Y | Y |  | electrochemical | OEM | Sensoric NO-3E100 | NO | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9947 | 1.32 | 1.001e-05 |  |  |  |  | linear | Rs=a0+a1*NO | O3 |  |  |  | NO2 | CO |  |  | REF | Sensor | [14 empty] | 0.23 ± 0.033 |  | 2.2 ± 0.23 | -0.16 ± 0.033 | 0.52 ± 0.86 | -53.4 ± 0.76 | -0.17 ± 0.052 | [7 empty] | 74.9 | ppb.min |  |  | 0-1000 | ppb | [10 empty] | YES |  | YES |  |  | 120 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.82 | [6 empty] | MLR | Model 0 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 1 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 2 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO-B4 | Y | Y |  | electrochemical | OEM | AlphaSense CO-B4 | NO | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 3 | [24 empty] | 0.01 |  | 0.01 | [32 empty] | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
CityTechnology |  |  | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | Sensoric NO2-3E50 | NO2 | ppb | Y | open source | LAB |  | 98 |  | 1 | hour | 0.9992 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.0016 | -0.013 | 1.5 |  | 0.16 | 0.062 | 0.8 | SD | [13 empty] | 2.3 |  | 1.2 | < -0.196 | [9 empty] | 100 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | NO2-3E50 | Y | Y | https://www.citytech.com/PDF-Datasheets/no23e50.pdf | electrochemical | OEM | Sensoric NO2-3E50 | NO2 | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9986 | 1.67 | -3.6e-05 |  |  |  |  | linear | Rs=a0+a1*NO2 | O3 |  |  | NO |  |  |  |  | REF | Sensor | [15 empty] | -0.032 ± 0.0032 | -1.22 ± 0.050 | 1.53 ± 0.015 | -2.48 ± 0.049 | 26.8 ± 0.3 | 0.56 ± 0.43 | [7 empty] | 16.7 | ppb.min |  |  | 0-50 | ppb | [10 empty] | YES |  | YES |  |  | 100 | EUR | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense NO2-B4 | NO2 | ppb | Y | open source | LAB |  | 80 |  | 1 | hour | 0.9997 | [8 empty] | O3 | [7 empty] | REF | Sensor | linear | [10 empty] | 22 | 60 |  |  | -0.032 | < -1.3e-03 | 1.5 |  | 0.47 | 0.13 | 2.9 | SD | [13 empty] | 11.5 |  | 3.2 | < -0.497 |  |  |  |  | YES |  | YES |  |  | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.108 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.009 | 522.71 | -73.28 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.23 | [6 empty] | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | linear | 0.002 | 991.24 | -139.52 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.351 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.026 | -108.36 | 17.07 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | NO2-B4 | Y | N |  | electrochemical | OEM | AlphaSense NO2-B4 | NO2 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.679 | [6 empty] | MLR | Rs=a1*NO2+a2*O3+a3*T+a4*RH+a0 | O3 | T | RH | [5 empty] | REF | Sensor | linear | 0.086 | -33.59 | 6.96 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM | AlphaSense NO2-B42F | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.84 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 6.5 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM | AlphaSense NO2-B42F | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.79 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 5.8 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM | AlphaSense NO2-B42F | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 7.5 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | NO2-B42F | Y | N |  | electrochemical | OEM | AlphaSense NO2-B42F | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.87 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [15 empty] | 0.01 | 0.01 | 1 |  |  |  | 5.8 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Sun et al. 2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1916 | NO2-B43F | Y | Y | http://www.alphasense.com/index.php/products/nitrogen-dioxide-2/ | electrochemical | OEM | AlphaSense NO2-B42F | NO2 | ppb | Y | open source | LAB |  |  |  | 1 | hour | [7 empty] | MLR | NO2 = a1*Rs +a2*T +a3*RH +a4 |  | T | RH | [5 empty] | Sensor | REF | linear | 0.99 | 0 | 0.95 | [45 empty] | 50 | EUR | https://www.mdpi.com/1424-8220/17/8/1916 |  | JRC-FK
CityTechnology |  |  | Gerboles et al. 2015(a) | 2015 | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | Sensoric O3-31F | O3 | V/ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.9992 | 1.617 | -0.0001858 |  |  |  |  | linear | Rs=a0+a1*O3 |  |  |  | NO | NO2 | CO |  |  | REF | Sensor | [14 empty] | 0.53 ± 0.026 | -0.018 ± 0.0030 | -0.11 ± 0.035 |  | -0.10 ± 0.017 | 6.6 ± 0.90 | 0.058 ± 0.55 | [7 empty] | 3.2 | ppb.min |  |  | 0-1000 | ppb | [10 empty] | YES |  | YES |  |  | 500 | USD | http://publications.jrc.ec.europa.eu/repository/bitstream/JRC97581/lbna27469enn.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et al. 2015 (a) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | LAB |  | 84 |  | 1 | hour | 0.9998 | [12 empty] | NO2 |  |  |  | REF | Sensor | linear | [13 empty] | 0.926 | -0.042 | -0.066 |  |  |  | 0.4 | 0.4 | SD | [13 empty] | 0.9 |  | 0.7 | < 0.016 | [9 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Published%20proceedings_corrected.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.023 | [6 empty] | linear | Rs=a1*O3+a0 | [8 empty] | REF | Sensor | linear | 0.212 | -142.99 | 4.83 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  | Euramet - MacPoll | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | 0.503 | [6 empty] | MLR | Rs=a1*O3+a2*NO2+a3*NO2+a4*H2O+a0 |  |  |  |  | NO2 |  |  |  | REF | Sensor | linear | 0.479 | -37.97 | 2.33 | [45 empty] | 50 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.7 | [6 empty] | MLR | Model 0 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 4.5 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.73 | [6 empty] | MLR | Model 1 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 6.1 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.73 | [6 empty] | MLR | Model 2 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 13.4 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
Alphasense |  |  | Wei et al. 2018 | 2018 | https://www.mdpi.com/1424-8220/18/2/59 | O3-B4 | Y | N |  | electrochemical | OEM | AlphaSense O3-B4 | O3 | ppb | Y | open source | LAB |  |  |  | 1 | hour | 0.72 | [6 empty] | MLR | Model 3 |  | T | RH | [5 empty] | Sensor | REF | [21 empty] | 7 | SD | [21 empty] | YES |  | YES |  |  | 50 | EUR | https://www.mdpi.com/1424-8220/18/2/59 |  | JRC-FK
CityTechnology |  | MacPoll.eu | Spinelle et at. 2015 Part A (b) | 2015 | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | O3-3E1F | Y | Y | https://www.citytech.com/PDF-Datasheets/o33e1f.pdf | electrochemical | OEM | City Technology - sensoric - O3-3E1F | O3 | ppb | Y | open source | FIELD | semi-rural |  |  | 1 | hour | [7 empty] | ANN | O3 =f(O3 3E1F, MICS-2710, TGS-5042) | [8 empty] | REF | Sensor | linear | 0.915 | -2.77 | 1.12 | [45 empty] | 500 | USD | http://db-airmontech.jrc.ec.europa.eu/download/Spinelle%20et%20al_2015_Field%20calibration%20of%20a%20cluster%20of%20low-cost%20available%20sensors%20for%20air%20quality.pdf | ANN with row data | JRC-FK
LiberaIntentio | Atmo-Normandie |  | Wastine_2019_a | 2019 | AirSensEUR-Atmo-Normandie | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 | µg/m3 | Y | open source | FIELD | urban |  | 60 | 1 | hour | [7 empty] | ANN | All parameters: NO, NO2, O3, CO sensors, T, RH, Press | [8 empty] | REF | Sensor | linear | 0.86 |  | 0.91 | 8.2 | [44 empty] | 1600 | EUR | AirSensEUR-Atmo-Normandie | winter field-test (10 min. data) | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.164 | 30.771 | 0.609 |  |  |  |  | 98 |  | 22 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.2519 | 24.712 | 0.6344 |  |  |  |  | 98 |  | 31 | [38 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/iqair-airvisual-pro---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
IQAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/iqair-airvisual-pro---lab-evaluation.pdf?sfvrsn=6 | AirVisual Pro | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | LAB |  |  | 10 | 5 | min | [17 empty] | Sensor | BAM | linear | 0.99 | 0.88 | 0.53 |  |  |  |  | 100 |  |  | 20 | 40 | [12 empty] | -6.2 | 1140% | % | [21 empty] | 270 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/iqair-airvisual-pro---lab-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt410 | electrochemical | sensor system | AlphaSense | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4868 | 0.0795 | -0.2051 |  |  |  |  | 85 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt411 | electrochemical | sensor system | AlphaSense | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4001 | 0.0744 | -0.2761 |  |  |  |  | 85 |  | 0.13 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt412 | electrochemical | sensor system | AlphaSense | O3 | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.5833 | 0.084 | -0.1855 |  |  |  |  | 85 |  | 0.24 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.486 | 1.955 | 0.0015 |  |  |  |  | 100 |  | 6100 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.436 | 1.8609 | 0.0034 |  |  |  |  | 100 |  | 2900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.502 | 1.69 | 0.0025 |  |  |  |  | 100 |  | 3900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.3953 | 6.71 | 0.0364 |  |  |  |  | 100 |  | 530 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5017 | 4.1087 | 0.093 |  |  |  |  | 100 |  | 210 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.147 | 23.896 | 0.0429 |  |  |  |  | 100 |  | 50 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.628 | 1.4255 | 0.0016 |  |  |  |  | 100 |  | 6100 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.54 | 1.1799 | 0.0037 |  |  |  |  | 100 |  | 2900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.605 | 1.3708 | 0.0026 |  |  |  |  | 100 |  | 3900 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.43 | 9.85 | 0.0303 |  |  |  |  | 100 |  | 530 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.588 | 6.245 | 0.0839 |  |  |  |  | 100 |  | 210 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
NuWave Sensors |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 | Cair | Y | Y | https://nuwavesensors.com/wp-content/uploads/2018/02/cair_Brochure.pdf | OPC | sensor system | NuWave Sensors | PM10-2.5 | #/ft3 | Y | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.3939 | 21.212 | 0.0867 |  |  |  |  | 100 |  | 50 | [38 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/cair---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.612 | 11.439 | 0.9303 |  |  |  |  | 100 |  | 8 | [38 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.55 | 41.991 | 1.064 |  |  |  |  | 100 |  | 3.8 | [38 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.609 | 12.972 | 0.7418 |  |  |  |  | 100 | [40 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Met One |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 | E-Sampler | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 300 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.552 | 15.878 | 0.8323 |  |  |  |  | 100 | [40 empty] | 5500 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/met-one-e-sampler---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9822 | 1.626 | 0.8125 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9849 | 2.0202 | 0.6981 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9784 | 1.4848 | 0.6969 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9655 | 2.3172 | 0.7873 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.855 | 1.8793 | 0.638 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9847 | 1.9355 | 0.8283 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9791 | 2.7288 | 0.6241 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9648 | 3.0028 | 0.5171 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
Dylos |  |  | Dacunto et al. 2015 | 2015 | https://www.ncbi.nlm.nih.gov/pubmed/26487426 | Dylos DC1700 | Y | Y | http://www.carbideprocessors.com/pages/woodworking/dylos-air-quality-monitor-exposure-limits.html | OPC | sensor system | Dylos DC1700 | PM2.5 | #/(ft3/100) | Y | open source | FIELD |  |  |  | 1 | min | 0.9 |  | 3.91e-06 | 1.803 |  |  |  | exp | PM2.5=a1*Rs² | [59 empty] | 475 | USD | https://www.ncbi.nlm.nih.gov/pubmed/26487426 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9785 | 2.4318 | 0.5655 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9386 | 3.4081 | 0.5829 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9028 | 2.5956 | 0.5302 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.9754 | 2.9308 | 0.636 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7037 | 7.3073 | 0.6882 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6785 | 7.7173 | 0.5325 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6967 | 6.9361 | 0.6099 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6677 | 8.0026 | 0.5985 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6598 | 7.0798 | 0.5871 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.76 | 1.76 | 0.22 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.53 | 2.71 | 0.21 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.81 | 0.32 | 0.23 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Mukherjee_et_al_2017 | 2017 | https://www.mdpi.com/1424-8220/17/8/1805 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | rural |  |  | 1 | hour | [7 empty] | linear | Rs=0.518+0.00274*count (hppcf) | [8 empty] | BAM | Sensor | linear | 0.67 | 4.62 | 0.16 |  |  |  |  | 100 | [40 empty] | 310 | USD | https://www.mdpi.com/1424-8220/17/8/1805 |  | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | 3.9 | 0.2 | [19 empty] | 6.6 | 6.6 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt Table 2 | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.94 | 0.7 | 0.03 | [19 empty] | 9.5 | 9.5 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | 5 | 1.7 | [19 empty] | 15 | 15 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | -15 | 0.5 | [19 empty] | 4.8 | 4.8 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.94 | -19 | 0.04 | [19 empty] | 8.7 | 8.7 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | 24 | 0.9 | [19 empty] | 14 | 14 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | -18 | 0.5 | [19 empty] | 4.2 | 4.2 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | salt | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.96 | -17 | 0.05 | [19 empty] | 6.1 | 6.1 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | welding fume | JRC-FK
Alphasense |  |  | Sousan et al. 2016 (a) | 2016 | http://dx.doi.org/10.1080/02786826.2016.1232859 | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | LAB |  |  |  | 3 | min | [17 empty] | GRIMM | Sensor | linear | 0.99 | 105 | 1.6 | [19 empty] | 16 | 16 | [15 empty] | YES | 3000 | NO | [6 empty] | 310 | USD | http://dx.doi.org/10.1080/02786826.2016.1232859 | Arizona dust | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.2025 | -0.3 | 0.4 | 1.08 |  |  |  | 76 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.115 | -1.66 | 0.49 |  |  |  |  | 71 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0121 | 0.6 | 0.07 |  |  |  |  | 82 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.2209 | 2.98 | 0.45 | 1.01 |  |  |  | 76 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.4624 | -1.06 | 0.54 |  |  |  |  | 71 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
AlphaSense |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | AlphaSense | PM10 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.04 | 2.86 | 0.12 |  |  |  |  | 82 | [40 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.86 | 0.02 | 3.19 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.91 | 0.01 | 2.63 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.63 | 0.04 | 3.53 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.76 | 0.03 | 3.02 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | TSI | Sensor | linear | 0.57 | 0.03 | 2.29 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [17 empty] | GRIMM | Sensor |  | 0.67 | 0.02 | 2.06 | [45 empty] | 310 | USD | https://www.atmos-meas-tech.net/11/709/2018/ |  | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor | linear | 0.84 | -3.7 | 1.12 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 3, correction of Kohler theory, averages, 250 POUNDS | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | GRIMM | Sensor | linear |  |  | 1.3 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 3, correction of Kohler theory, averages, 250 POUNDS | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor | linear |  |  | 0.905 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 3, correction of Kohler theory, averages, 250 POUNDS | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | GRIMM | Sensor | linear |  |  | 1.295 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 3, correction of Kohler theory, averages, 250 POUNDS | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | FIELD | traffic |  | 15 | 1 | min | [11 empty] | RH | [5 empty] | TEOM | Sensor | linear | 0.42 |  | 5.25 | 19.84 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 5 | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | black box | FIELD | traffic |  | 15 | 1 | min | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor | linear | 0.73 |  | 1.15 | 1.66 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 5 | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | Y | open source | FIELD | traffic |  | 15 | 1 | min | [7 empty] | Kholer | non linear model including RH per bin |  |  | RH | [5 empty] | TEOM | Sensor | linear | 0.75 |  | 1.05 | 1.37 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 5 | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | traffic |  | 15 | 1 | min | [11 empty] | RH | [5 empty] | TEOM | Sensor | linear | 0.34 |  | 4.59 | 35.78 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 6 | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | traffic |  | 15 | 1 | min | [7 empty] | Kholer | non linear model including RH |  |  | RH | [5 empty] | TEOM | Sensor | linear | 0.75 |  | 1.43 | 5.91 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 6 | JRC-FK
Alphasense |  |  | Di-Antonio_developing_2018 | 2018 | https://www.mdpi.com/1424-8220/18/9/2790/pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | open source | FIELD | traffic |  | 15 | 1 | min | [7 empty] | Kholer | non linear model including RH per bin |  |  | RH | [5 empty] | TEOM | Sensor | linear | 0.78 |  | 1.01 | 3.74 | [44 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | table 6 | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [11 empty] | RH | [5 empty] | TEOM | Sensor | linear | 0.71 |  | 3.2 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 2, All RH, median | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [11 empty] | RH | [5 empty] | GRIMM | Sensor | linear | 0.715 |  | 3.75 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 2, All RH, median | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [11 empty] | RH | [5 empty] | TEOM | Sensor | linear | 0.655 |  | 2.7 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 2, All RH, median | JRC-FK
Alphasense |  |  | Crilley et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/709/2018/ | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | Y | black box | FIELD | urban |  | 15 | 1 | hour | [11 empty] | RH | [5 empty] | GRIMM | Sensor | linear | 0.67 |  | 3.65 | [45 empty] | 320 | EUR | https://www.atmos-meas-tech.net/11/709/2018/ | Results of Table 2, All RH, median | JRC-FK
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.3433 |  |  |  |  | 114.6 | [17 empty] | 17.6 | [24 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.353 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.3766 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Alphasense |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.556 |  |  |  |  | 114.6 | [42 empty] | 310 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7014 | 7.4581 | 0.688 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9859 | 1.5968 | 0.8152 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9891 | 1.997 | 0.6996 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9818 | 1.4361 | 0.7 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9679 | 2.3088 | 0.7906 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9063 | 1.3309 | 0.6764 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.989 | 1.9103 | 0.8307 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9831 | 2.7136 | 0.6255 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9693 | 2.9952 | 0.5177 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9821 | 2.405 | 0.5678 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9416 | 3.4198 | 0.5849 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9384 | 2.2195 | 0.5511 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9801 | 2.9112 | 0.6376 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7212 | 7.3213 | 0.6907 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6966 | 7.7174 | 0.5346 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7065 | 7.0549 | 0.5987 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6851 | 8.0766 | 0.6008 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.699 | 6.6847 | 0.6082 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7235 | 7.4145 | 0.6922 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9928 | 1.5014 | 0.8266 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9918 | 2.0869 | 0.6964 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.99 | 1.1948 | 0.7185 |  |  |  |  | 95 | 98 | 12 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9745 | 2.529 | 0.7749 |  |  |  |  | 95 | 98 | 10 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9829 | 0.4563 | 0.7402 |  |  |  |  | 95 | 98 | 13 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9946 | 2.0312 | 0.8237 |  |  |  |  | 95 | 98 | 11 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9873 | 2.721 | 0.6305 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9716 | 3.1451 | 0.5152 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9873 | 2.2258 | 0.5808 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9489 | 3.7076 | 0.5756 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9859 | 1.7418 | 0.5812 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9854 | 3.0464 | 0.6345 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7731 | 6.9258 | 0.73 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7482 | 7.5023 | 0.559 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7669 | 6.3589 | 0.6559 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7338 | 8.0511 | 0.6212 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7833 | 5.7248 | 0.66886 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7796 | 7.1338 | 0.7249 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9194 | 2.3114 | 0.613 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8967 | 2.6344 | 0.5046 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9208 | 1.9955 | 0.5569 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8682 | 3.0635 | 0.5693 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.8791 | 1.8201 | 0.5404 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.9162 | 2.5066 | 0.6248 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.654 | 11.207 | 0.7476 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6247 | 11.701 | 0.5753 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6505 | 10.81 | 0.6641 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6067 | 12.161 | 0.6427 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6379 | 10.485 | 0.6598 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.6536 | 11.332 | 0.7479 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9713 | 2.4968 | 0.6057 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9528 | 2.919 | 0.4941 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9721 | 2.0167 | 0.5582 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9269 | 3.4751 | 0.551 |  |  |  |  | 95 | 98 | 15 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.9657 | 1.5789 | 0.5571 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6092 | 2.8139 | 0.6092 |  |  |  |  | 95 | 98 | 17 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7336 | 11.161 | 0.7717 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7017 | 11.844 | 0.5873 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7313 | 10.527 | 0.6949 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6827 | 12.469 | 0.65 |  |  |  |  | 95 | 98 | 19 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7475 | 9.85 | 0.7086 |  |  |  |  | 95 | 98 | 20 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7359 | 11.415 | 0.7641 |  |  |  |  | 95 | 98 | 18 | [38 empty] | 200 | USD | http://db-airmontech.jrc.ec.europa.eu/download/purple-air-pa-ii---field-evaluation.pdf |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 1 | -8.49 | 1.49 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 85-95 over 0-175 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -15.64 | 1.31 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 54-96 over 0-250 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 | PA-II | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | LAB |  |  | 35 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.94 | -19.79 | 2.38 |  |  |  |  | 95.9 | 96.7 |  | 20 | 40 | [12 empty] | 44-73 over 21-180 ug/m3 | [23 empty] | 200 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/purple-air-pa-ii---lab-evaluation.pdf?sfvrsn=4 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7133 | 4.2246 | 0.4975 |  |  |  |  | 80 | 100 | 14 | [14 empty] | 35- 69 over 16-270 ug/m3 | [23 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7277 | 3.9572 | 0.4781 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7425 | 2.8438 | 0.566 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7761 | 4.4018 | 0.4855 |  |  |  |  | 80 | 100 | 14 | [14 empty] | 35- 69 over 16-270 ug/m3 | [23 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7943 | 4.0539 | 0.4727 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 30 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8276 | 2.8625 | 0.5641 |  |  |  |  | 80 | 100 | 15 | [38 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/sainsmart---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
SainSmart |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/sainsmart---lab-evaluation.pdf?sfvrsn=6 | Pure Morning P3 | Y | Y |  | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | LAB |  |  | 30 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.99 | -4.13 | 0.79 |  |  |  |  | 99.6 | 100 |  | 20 | 40 | [36 empty] | 170 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/laboratory-evaluations/sainsmart---lab-evaluation.pdf?sfvrsn=6 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.008 | 31.649 | 0.0576 |  |  |  |  | 95 |  | 31 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0006 | 19.311 | 0.0457 |  |  |  |  | 96 |  | 20 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y |  | nephelometer | sensor system | Shinyei PPD42NS | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0002 | 19.786 | 0.0191 |  |  |  |  | 88 |  | 20 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.4326 | 15.048 | 0.6426 |  |  |  |  | 95 |  | 10 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7175 | 13.507 | 0.2062 |  |  |  |  | 96 |  | 30 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
uHoo |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 | uHoo | Y | Y | https://seetheair.wordpress.com/2017/09/04/review-uhoo-aqi-monitor/ | MOs | sensor system | SGX Sensortech MICS-2614 | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7156 | 5.8943 | 0.6287 |  |  |  |  | 88 |  | 18 | [38 empty] | 300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/uhoo---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.2832 | 0.2416 | 0.6047 |  |  |  |  | 71 |  | 0.1 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.3136 | 0.2377 | 0.7756 |  |  |  |  | 71 |  | 0.07 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.2867 | 0.2433 | 0.572 |  |  |  |  | 71 |  | 0.12 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.00743 | 0.0104 | -0.0166 |  |  |  |  | 85 |  | 0.15 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.00717 | 0.0104 | -0.0205 |  |  |  |  | 85 |  | 0.14 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 | Vaisala AQT410 v.1.11 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.0821 | 0.0105 | -0.0142 |  |  |  |  | 85 |  | 0.16 | [38 empty] | 3700 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala---field-evaluation.pdf?sfvrsn=14 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8026 | 0.233 | 0.9572 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8219 | 0.2417 | 0.8902 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.8297 | 0.2316 | 1.0079 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.85 | 0.2315 | 0.9802 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8721 | 0.2377 | 0.9134 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.8746 | 0.2305 | 1.0267 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8367 | 0.2432 | 0.9015 |  |  |  |  | 96 |  | 0.2 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.8166 | 0.2491 | 0.8211 |  |  |  |  | 96 |  | 0.21 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | CO | ppm | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.86123 | 0.2504 | 0.9105 |  |  |  |  | 96 |  | 0.19 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.5447 | -0.319 | 0.5114 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.4372 | 1.6073 | 0.3927 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | FRM | linear | 0.6051 | -1.7053 | 0.5 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.5643 | -0.3845 | 0.5203 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.4497 | 1.6573 | 0.3988 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | FRM | linear | 0.6225 | -1.8693 | 0.508 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.4275 | 4.3265 | 0.3246 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.2954 | 6.318 | 0.2212 |  |  |  |  | 96 |  | 29 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2017 | http://www.aqmd.gov/aq-spec/evaluations/summary | Vaisala AQT410 v.1.15 | Y | Y |  | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | FRM | linear | 0.4973 | 2.5653 | 0.345 |  |  |  |  | 96 |  | 29.5 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8656 | -0.0595 | 0.8401 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.868 | -0.0514 | 0.8563 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8755 | 0.0746 | 0.8294 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8423 | 3.4367 | 0.5649 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8543 | 3.513 | 0.589 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.841 | 3.5134 | 0.551 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1235 | 24.687 | 0.5098 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1304 | 26.38 | 0.551 |  |  |  |  | 99.8 |  | 20 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.1283 | 23.812 | 0.517 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8686 | -0.1195 | 0.8454 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8736 | -0.1482 | 0.8651 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8787 | 0.0161 | 0.8346 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8563 | 3.4088 | 0.5661 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8703 | 3.467 | 0.5914 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8544 | 3.4895 | 0.5522 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8693 | 0.3665 | 0.8092 |  |  |  |  | 99.8 |  | 12 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/egg | OPC | sensor system | Plantower - PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8313 | 0.6412 | 0.8136 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.8859 | 0.2977 | 0.8199 |  |  |  |  | 99.8 |  | 11 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.7575 | 4.311 | 0.518 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.7466 | 4.6687 | 0.5327 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.7915 | 3.9654 | 0.5335 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.5929 | 5.3928 | 0.5294 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6356 | 5.5364 | 0.5768 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.5769 | 5.3565 | 0.5157 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6246 | 5.5996 | 0.519 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6447 | 6.0854 | 0.5413 |  |  |  |  | 99.8 |  | 17 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6598 | 5.0593 | 0.5428 |  |  |  |  | 99.8 |  | 18 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1134 | 32.337 | 0.53 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1156 | 33.645 | 0.5179 |  |  |  |  | 99.8 |  | 20 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
Air Quality Egg |  | Air Quality Egg | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | Egg (2018) | Y | Y | https://airqualityegg.com/home | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.1121 | 31.614 | 0.5258 |  |  |  |  | 99.8 |  | 21 | [38 empty] | 249 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/air-quality-egg---2018-model---field-evaluation.pdf?sfvrsn=12 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7041 | 2.1213 | 0.7802 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6825 | 1.7206 | 0.799 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.6854 | 1.499 | 0.8223 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7348 | 1.9365 | 0.7988 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7119 | 1.523 | 0.8178 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7166 | 1.2854 | 0.8434 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.9058 | 1.5949 | 0.8316 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8705 | 1.2561 | 0.8386 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM1 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.881 | 0.9495 | 0.8746 | [6 empty] | 10 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5649 | 4.9727 | 0.5269 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5496 | 5.1093 | 0.5147 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.5552 | 4.9914 | 0.5287 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.612 | 4.7844 | 0.5383 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5836 | 4.9388 | 0.5247 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5901 | 4.8162 | 0.5395 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7725 | 4.1765 | 0.5413 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7451 | 4.9157 | 0.5226 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7416 | 4.88 | 0.534 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5485 | 5.6826 | 0.5942 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5184 | 5.9531 | 0.5724 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.5135 | 0.59431 | 0.5811 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.5714 | 6.2735 | 0.5421 | [6 empty] | 15.5 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.534 | 6.6015 | 0.515 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.5228 | 6.6436 | 0.5216 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0507 | 26.368 | 0.5531 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0401 | 27.43 | 0.4931 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.0388 | 27.487 | 0.4914 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0602 | 26.139 | 0.5639 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0475 | 27.255 | 0.5016 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0459 | 27.306 | 0.5 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0655 | 27.323 | 0.4978 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0569 | 27.909 | 0.4621 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.0488 | 27.445 | 0.4353 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0672 | 32.293 | 0.6137 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0529 | 36.496 | 0.5457 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.0491 | 36.752 | 0.533 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.1033 | 34.746 | 0.643 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0896 | 35.494 | 0.5971 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
AirThinx |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 | AirThinx | Y | Y | https://airthinx.io/products/iaq/ | OPC | sensor system | PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.0811 | 35.901 | 0.5777 | [6 empty] | 18 | [38 empty] | 1000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/airthinx-iaq--field-evaluation.pdf?sfvrsn=18 |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt413 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.6653 | 9.899 | 0.7531 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt414 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.7077 | 6.6974 | 0.8982 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt415 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.8118 | 2.5717 | 0.9888 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt416 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.6679 | 9.014 | 0.7637 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt417 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.7145 | 5.6011 | 0.9164 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt418 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.8148 | 2.1074 | 0.9954 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt419 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.4434 | 22.405 | 0.3863 |  |  |  |  | 96 |  | 38 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6161 | 17.796 | 0.5344 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Vaisala |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/vaisala-aqt410-v-1-15---field-evaluation.pdf?sfvrsn=6 | AQT-410 v.1.11 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt421 | electrochemical | sensor system | AlphaSense | O3 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [17 empty] | Sensor | REF | linear | 0.6853 | 14.53 | 0.623 |  |  |  |  | 96 |  | 39 | [38 empty] | 3700 | USD | http://www.aqmd.gov/aq-spec/evaluations/summary |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8669 | 4.1665 | 0.9373 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8412 | 4.9107 | 0.9263 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.7842 | 3.5351 | 0.9539 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8481 | 4.0137 | 0.9872 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8745 | 4.1148 | 0.9418 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8484 | 4.815 | 0.9318 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8042 | 3.4997 | 0.9622 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8562 | 3.9516 | 0.9928 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9225 | 3.6646 | 0.9773 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9126 | 4.819 | 0.9559 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.8645 | 3.6032 | 0.9468 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | GRIMM | linear | 0.9233 | 3.5683 | 1.0555 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8435 | 4.0564 | 0.9077 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8302 | 3.7918 | 0.9194 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.533 | 3.4347 | 1.3074 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8372 | 4.0173 | 0.9689 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3/mV | N | black box | LAB |  |  |  | 15 | min | [17 empty] | Sensor | REF | linear | 0.99 | 266.33 | 9.39 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ |  | JRC-FK
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | indoor |  |  | 48 | hour | [17 empty] | Sensor | GRAVIMETRIC | linear | 0.90253 |  | 1.4969 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ |  | JRC-FK
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3/mV | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.9 | 12 | 2.78 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ |  | JRC-FK
Berkeley Air |  |  | Pillarisetti et al 2017 | 2017 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ | PATS+ | Y | Y |  | OPC | sensor system |  | PM2.5 | µg/m3/mV | N | black box | FIELD | indoor |  |  | 1 | hour | [17 empty] | Sensor | DustTrak | linear | 0.96 | -0.05 | 0.92 | [45 empty] | 500 | USD | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5579926/ |  | JRC-FK
Platypus |  |  | Williams et al. 2014(c) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf | Platypus NO2 | N | Y |  | MOs | sensor system |  | NO2 | ppb | N | black box | LAB |  |  |  | 1 | hour | [20 empty] | 0.7956 | [32 empty] | 18 | ppb | [13 empty] | 50 | USD | http://db-airmontech.jrc.ec.europa.eu/download/SENSOR.pdf |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.9247 | 3.9097 | 0.9111 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.9318 | 3.8435 | 0.9201 |  |  |  |  | 90 |  | 15 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.5754 | 3.5293 | 1.1949 |  |  |  |  | 90 |  | 5 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | OPC | sensor system | Nova Fitness SDS 011 | PM2.5 | µg/m3 | Unknown, likely unavaialble | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | AQY PM2.5 corrected with AQY RH in real-time |  |  | RH | [5 empty] | Sensor | BAM | linear | 0.8849 | 4.4864 | 0.8085 |  |  |  |  | 90 |  | 10 | [14 empty] | ±10 | 5% | µg/m3 | 1 | µg/m3 |  |  | 0-1000 | µg/m3 | [15 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7477 | -1.7229 | 1.1845 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7831 | -0.2971 | 0.9549 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.9153 | 4.5088 | 0.9153 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.9801 | 2.0031 | 0.9801 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.7787 | -2.2543 | 1.2225 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8063 | 0.8664 | 0.978 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.2058 | 2.6131 | 1.1443 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8144 | 1.9001 | 0.9913 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8313 | -1.4757 | 1.1741 |  |  |  |  | 98 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8155 | -3.8834 | 1.0972 |  |  |  |  | 99 |  | 25 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.3971 | -1.1867 | 1.5499 |  |  |  |  | 97 |  | 10 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/product/aqy-micro-air-quality-station | electrochemical | sensor system | AlphaSense | NO2 | ppb | N | black box | FIELD | rural |  | 60 | 24 | hour | [8 empty] | Calculation by new on-instrument Aeroqual algorithm (AQ-SPEC version 2) | O3 |  | RH | [5 empty] | Sensor | FRM | linear | 0.8324 | 3.7397 | 0.8904 |  |  |  |  | 99 |  | 18 | [14 empty] | 8 | 8% | ppb | 2 | ppb |  |  | 0-500 | ppb | [11 empty] | YES |  |  |  | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Shinyei |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.0033 | 5.2302 | 0.0659 | [8 empty] | > 95% (spike) | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Shinyei |  |  | Williams et al. 2014(b) | 2014 | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 1 | 5 | min | [17 empty] | GRIMM | Sensor | linear | 0.1516 | 2.756 | 0.2917 | [8 empty] | > 95% (spike) | [36 empty] | 1000 | USD | http://db-airmontech.jrc.ec.europa.eu/download/PMSENSOREVALUATIONREPORTEPA-600-R-14-464.PDF |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5041 | 0.24 | 0.58 | 0.2 |  |  |  | 73 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.5184 | 0.8 | 0.54 |  |  |  |  | 87 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Shinyei |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | PMS-SYS-1 | Y | Y |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.0001 | 4.35 | 0.42 |  |  |  |  | 92 | [40 empty] | 1000 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 10 | hour | [17 empty] | REF | Sensor | linear | 0.84 | [29 empty] | 3.22 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.845 | 5.22 | 1.26 |  |  |  |  | linear | Rs=a0+a1*PM2.5 | [8 empty] | TEOM | Sensor | [31 empty] | 10.5 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9514 | -1.4532 | 1.438 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.4 | 3.7 | 0.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 2.01 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.41 | 4.5 | 0.7 | -0.06 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 2.07 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 6 | hour | 0.8 | -1.9 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.53 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.84 | -2.4 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.35 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 24 | hour | 0.93 | -4.2 | 1.6 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 0.15 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | min | 0.66 | -2.1 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 27 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | min | 0.93 | -3.5 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | min | 0.94 | -1.5 | 1.5 | -0.09 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.66 | -2.1 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 27 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.93 | -3.5 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 10 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 1 | hour | 0.95 | -1.4 | 1.5 | -0.09 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 8 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 6 | hour | 0.73 | -2.4 | 1.8 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 22 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 6 | hour | 0.92 | -3.2 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 10 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 6 | hour | 0.95 | -0.3 | 1.5 | -0.13 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 8 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.84 | -3.4 | 1.9 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 15 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 12 | hour | 0.93 | 8.7 | 2.2 | -0.49 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | T640 | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | suburban |  |  | 24 | hour | 0.94 | -4.1 | 2 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | T640 | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | 0.61 | 5.1 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 46 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | 0.75 | 74 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 35 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | 0.74 | 74 | 1.4 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 24 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | 0.61 | 88 | 1.2 | -2.3 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 46 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | 0.78 | 276 | 1.1 | -7.9 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 30 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 6 | hour | 0.8 | -5.8 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 32 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 6 | hour | 0.87 | 65 | 1.4 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 25 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 6 | hour | 0.86 | 65 | 1.4 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 16 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 6 | hour | 0.81 | 47 | 1.6 | -1.4 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 30 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 6 | hour | 0.9 | 248 | 1.2 | -7 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 18 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 12 | hour | 0.84 | -6.5 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 18 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 12 | hour | 0.88 | 74 | 1.3 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 19 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 12 | hour | 0.86 | 74 | 1.3 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 12 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 12 | hour | 0.89 | 330 | 1 | -10 |  |  |  | MLR | Rs=(a1*PM2.5+a2*T+a0)/(a+b*(RH2)/(1-RH)) |  | T | RH | [5 empty] | BAM | Sensor |  |  |  |  | 17 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | 1 | -4.5 | 1.7 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 17 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | 0.99 | 82 | 1.2 |  |  |  |  | linear | Rs=(a1*PM2.5+a0)/(a+b*(RH2)/(1-RH)) |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 11 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Zheng et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html | PMS3003 | Y | Y |  | OPC | OEM | Plantower | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | 1 | 82 | 1.2 |  |  |  |  | quad | Rs=(2a2(PM2.5-a1)2-a12+4a2a0)/4a2 |  |  | RH | [5 empty] | BAM | Sensor |  |  |  |  | 9 | [44 empty] | 30 | USD | https://www.atmos-meas-tech.net/11/4823/2018/amt-11-4823-2018-discussion.html |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS3003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 10 | hour | [17 empty] | REF | Sensor | linear | 0.85 | [47 empty] | 30 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Plantower |  |  | Laquai_particle_2017 | 2017 | https://www.researchgate.net/publication/320555036_Particle_Distribution_Dependent_Inaccuracy_of_the_Plantower_PMS5003_low-cost_PM-sensor | PMS5003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [49 empty] | YES | [6 empty] | 15 | EUR | https://www.researchgate.net/publication/320555036_Particle_Distribution_Dependent_Inaccuracy_of_the_Plantower_PMS5003_low-cost_PM-sensor |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.676 |  |  |  |  | 169.3 | [17 empty] | 6.98 | [24 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.71 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.75 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Plantower |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | PMS7003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.84 |  |  |  |  | 169.3 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Shinyei |  |  | Wang et al. 2015 | 2015 | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 | PPD42NS | Y | Y |  | nephelometer | OEM | PPD42NS | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2.5 | hour | 0.9452 | -0.353 | 33.6 |  |  |  |  | linear | [9 empty] | SidePak | Sensor | [23 empty] | 10 | 20 | [5 empty] | 6.44 | 4.59 | [8 empty] | YES | 300 | YES | 600 | YES | 400 | YES | 0.15 |  | 15 | USD | https://www.tandfonline.com/doi/full/10.1080/02786826.2015.1100710 |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | hour | 0.58 | [6 empty] | linear | PM2.5=a0+a1*Rs | [8 empty] | BAM | Sensor | [44 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 1 | hour | 0.59 | [6 empty] | MLR | PM2.5=a0+a1*Rs+a2*RH |  |  | RH | [5 empty] | BAM | Sensor | [44 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.92 |  |  | 2.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.94 |  |  | 2 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM3 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.94 |  |  | 2.1 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.9 |  |  | 1.7 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.93 |  |  | 1.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | GRIMM | Sensor | linear | 0.93 |  |  | 1.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.64 |  |  | 5.1 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.68 |  |  | 4.8 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | DustTrak | Sensor | linear | 0.7 |  |  | 4.6 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.56 |  |  | 3.5 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.61 |  |  | 3.4 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Shinyei |  |  | Holstius et al 2014 | 2014 | https://www.atmos-meas-tech.net/7/1121/2014/ | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | BAM | Sensor | linear | 0.58 |  |  | 3.6 | [39 empty] | YES |  | NO |  |  | 15 | USD | https://www.atmos-meas-tech.net/7/1121/2014/ |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9684 | -1.2815 | 1.2343 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Shinyei |  |  | Gao et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115000160 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | urban |  |  | 1 | min | [7 empty] | quad |  |  | T | RH | [5 empty] | DustTrak | Sensor | linear | 0.925 | [42 empty] | YES |  | YES |  |  | 15 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115000160 |  | JRC-FK
Shinyei |  |  | Gao et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S0269749115000160 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | FIELD | urban |  |  | 1 | hour | [7 empty] | quad |  |  | T | RH | [5 empty] | DustTrak | Sensor | linear | 0.885 | [42 empty] | YES |  | YES |  |  | 15 | USD | https://www.sciencedirect.com/science/article/pii/S0269749115000160 |  | JRC-FK
Shinyei |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | N | open source | LAB |  |  |  | 10 | min | [17 empty] | REF | Sensor | linear | 0.65 | [47 empty] | 15 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9492 | -22.029 | 1.5153 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 5 | min | [17 empty] | Sensor | REF | linear | 0.9635 | -6.6195 | 1.2611 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.961 | -1.4771 | 1.4403 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9761 | -1.0974 | 1.2298 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9612 | -22.026 | 1.5157 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 1 | hour | [17 empty] | Sensor | REF | linear | 0.9678 | -6.4263 | 1.2545 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.957 | 1.1417 | -0.8297 |  |  |  |  | 92 |  | 16 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | LAB |  |  |  | 1 | hour | 0.92 | 115 | 0.72 |  |  |  |  | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense CO-B4 | CO | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.92 | 26 | 0.89 |  |  |  |  | MLR | Rs=a1*CO+a2*T+a3*RH |  | T | RH | [5 empty] | REF | Sensor | [14 empty] | 0.085 |  |  | -0.053 | -0.034 | [31 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ | cross sensitivity come from Lewis et al. 2016 | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9653 | 1.2006 | -0.4091 |  |  |  |  | 76 |  | 18 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | LAB |  |  |  | 1 | hour | 0.31 | 34.8 | 0.3 |  |  |  |  | linear | Rs=a1*NO2+a0 | [8 empty] | REF | Sensor | [21 empty] | 3.5 ppb | [28 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
Carnegie Mellon University and |  |  | Zimmerman_et_al_2018 | 2018 | https://www.atmos-meas-tech.net/11/291/2018/ | RAMP | N | Y |  | electrochemical | sensor system | Alphasense NO2-B43F | NO2 | ppb | N | open source | FIELD | urban |  |  | 1 | hour | 0.27 | 8.8 | 0.33 |  |  |  |  | MLR | Rs=a1*NO2+a2*T+a3*RH |  | T | RH | [5 empty] | REF | Sensor | [50 empty] | USD | https://www.atmos-meas-tech.net/11/291/2018/ |  | JRC-FK
Plantower |  |  | Kelly et al 2017 | 2017 | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub | PMS1003 | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  |  | 1 | hour | 0.92 | -7.16 | -0.033 | 90.9 |  |  |  | exp | Rs= a0+a1*(1-exp(a2*PM2.5)) | [8 empty] | TEOM | Sensor | [31 empty] | 10.5 | [17 empty] | 20 | USD | https://www.sciencedirect.com/science/article/pii/S026974911632718X?via%3Dihub |  | JRC-FK
Aeroqual |  |  | Lin et al. 2015 | 2015 | https://www.sciencedirect.com/science/article/pii/S1352231014008607 | S-500 | Y | Y |  | MOs | sensor system | Aeroqual S-500 | NO2 | ppb | Y | black box | FIELD | urban |  |  | 1 | hour | [7 empty] | linear | Rs=NO2+a1*O3+a0 | O3 | [7 empty] | REF | Sensor | linear | 0.88 | -1.1 | 1.03 | [45 empty] | 500 | USD | https://www.sciencedirect.com/science/article/pii/S1352231014008607 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9612 | 1.4787 | -20.639 |  |  |  |  | 97 |  | 31 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Aeroqual |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 | AQY v0.5 | Y | updated | https://www.aeroqual.com/ | MOs | sensor system |  | O3 | ppb | N | black box | FIELD | rural |  | 60 | 8 | hour | [17 empty] | Sensor | REF | linear | 0.9641 | 1.2035 | -4.7653 |  |  |  |  | 100 |  | 21 | [38 empty] | 3000 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/aeroqual-aqy-v0-5---field-evaluation.pdf?sfvrsn=10 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.732 | 2.885 | 0.726 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7579 | 1.99 | 0.8029 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7384 | 2.9473 | 0.7151 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.846 | 2.91 | 0.711 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8705 | 1.976 | 0.791 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Clarity Movement Co |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 | Clarity Node | Y | Y | https://clarity.io/solution | OPC | sensor system | Sensirion-SPS30 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 120-240 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.852 | 2.905 | 0.7066 |  |  |  |  | 97 | 100 | 12 | [38 empty] | 1300 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/clarity-node---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.788 | 4.8544 | 1.703 |  |  |  |  | 100 | 100 | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.7718 | 4.9037 | 2.0427 |  |  |  |  | 100 | 100 | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Aeroqual |  |  | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf | S-500 | Y | Y |  | MOs | sensor system |  | NO2 | ppb | Y | black box | LAB |  | < 60 |  | 1 | hour | 0.8008 | -2.2563 | 0.8899 | [14 empty] | REF | Sensor | linear | [25 empty] | 15% |  | 1 | ppb |  |  | 0-0.2 | ppm | [15 empty] | 500 | USD | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | -0.69 | 0.66 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | 0.48 | 0.48 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Nova Fitness |  |  | Budde et al. 2018 | 2018 | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM10 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [11 empty] | RH | [5 empty] | SMPS | Sensor | linear |  | -0.53 | 0.65 | [38 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/325416608_Suitability_of_the_Low-Cost_SDS011_Particle_Sensor_for_Urban_PM-Monitoring/figures |  | JRC-FK
Shinyei |  |  | Austin et al. 2015 | 2015 | http://dx.plos.org/10.1371/journal.pone.0137789 | PPD42NS | Y | Y |  | nephelometer | OEM |  | PM2.5 | µg/m3 | Y | open source | LAB |  |  |  | 2 | min | [7 empty] | exp | Rs=a0(1-exp(-K*PM2.5)) | [8 empty] | APS | Sensor | linear | 0.97 |  | 5.4 | [27 empty] | 1 | [10 empty] | YES | [6 empty] | 15 | USD | http://dx.plos.org/10.1371/journal.pone.0137789 | Arizona Test Dust | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban | 10 |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.61 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban | 10 |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.64 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban | 10 |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.67 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova Fitness |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban | 10 |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.773 |  |  |  |  | 69 | [17 empty] | 6.9 | [24 empty] | 30 | EUR | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.71 | 0.965 | 0.891 | 4.21 | 2.97 | 0.04 | [14 empty] | 4.21 | SD | 9.64 |  |  | 98.16 | % | [21 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.68 | 2.97 | 0.816 | 4.49 | 3.18 | -1.25 | [14 empty] | 4.31 | SD | 9.64 |  |  | 86.82 | % | [21 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.55 | 0.867 | 0.708 | 5.45 | 3.84 | 1.87 | [14 empty] | 5.12 | SD | 9.64 |  |  | 80.76 | % |  |  |  |  | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.75 | 2.25 | 0.77 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.7 | 2.61 | 0.736 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | MLR | [9 empty] | TEOM | Sensor | linear | 0.6 | 3.57 | 0.635 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova Fitness |  |  | Laquai_particle_2017 | 2017 | https://www.researchgate.net/publication/322628807_Development_of_a_Calibration_Methodology_for_the_SDS011_Low-Cost_PM-Sensor_with_respect_to_Professional_Reference_Instrumentation | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour |  | 1.2203 | -0.509 |  |  |  |  | log | RsPM2.5/PM2.5=a1*ln(RSPM10/RSPM2.5)+a0 |  |  | RH | [28 empty] | 16 | SD | [19 empty] | YES |  | YES |  |  |  |  | 30 | EUR | https://www.researchgate.net/publication/322628807_Development_of_a_Calibration_Methodology_for_the_SDS011_Low-Cost_PM-Sensor_with_respect_to_Professional_Reference_Instrumentation |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.8 | 1.67 | 0.824 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.79 | 1.76 | 0.82 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | GRIMM | linear | 0.8113 | 4.375 | 1.9506 |  |  |  |  | 100 | 100 | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.798 | 4.7902 | 1.718 |  |  |  |  | 100 | 100 | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7811 | 4.8434 | 2.0593 |  |  |  |  | 100 | 100 | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8213 | 4.325 | 1.9628 |  |  |  |  | 100 | 100 | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8903 | 3.9835 | 1.9068 | [6 empty] | 3.9 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8767 | 4.1872 | 2.2491 | [6 empty] | 3 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM1 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.8991 | 4.033 | 2.0394 | [6 empty] | 3.8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6176 | 7.8671 | 0.8689 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6053 | 7.9458 | 1.027 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6911 | 6.4735 | 1.1354 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6948 | 5.9767 | 1.0526 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.6916 | 6.1498 | 1.2367 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.7619 | 5.2231 | 1.2741 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4064 | 9.638 | 0.6573 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4232 | 9.4925 | 0.8007 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4801 | 8.3687 | 0.882 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7496 | 3.4146 | 1.1942 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7852 | 3.5272 | 1.4267 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8143 | 3.0113 | 1.415 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
TSI |  |  | Karagulian et al. 2012 | 2012 | https://pubs.rsc.org/en/content/articlelanding/2012/em/c2em30099k#!divAbstract | SidePak AM510 | Y | N |  | nephelometer | sensor system |  | PM2.5 | µg/m3 | Y | open source | FIELD | rural |  |  | 30 | min | [17 empty] | TEOM-FDMS | Sensor | linear | 0.94 | 0 | 1 | [36 empty] | YES |  | YES |  | YES |  |  |  |  | 3000 | USD | https://pubs.rsc.org/en/content/articlelanding/2012/em/c2em30099k#!divAbstract |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.86 | -0.004 | 0.56 |  |  |  |  | 82 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.846 | -0.004 | 0.58 |  |  |  |  | 73 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Aeroqual |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | SM50 | Y | Y |  | MOs | OEM |  | O3 | ppb | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.9216 | -0.004 | 0.77 |  |  |  |  | 81 | [40 empty] | 500 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.523 | 9.3879 | 0.7119 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.5373 | 9.3947 | 0.8577 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.6116 | 8.3596 | 0.939 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.535 | 9.2966 | 0.7225 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.5511 | 9.2799 | 0.8734 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.6243 | 8.2691 | 0.9507 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7928 | 5.3506 | 1.141 | [6 empty] | 9.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.8103 | 5.5877 | 1.3466 | [6 empty] | 7.5 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.8464 | 5.0612 | 1.3403 | [6 empty] | 8 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Figaro |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | MOs | OEM | Figaro TGS-5042 | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.3 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.136 | 2.24 | 13.1 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | MOs | OEM | Figaro TGS-5042 | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.17 | [6 empty] | linear | Rs=a1*CO+a0 | [8 empty] | REF | Sensor | linear | 0.151 | 3.39 | 17.7 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Nova fitness | NILU |  | Liu et al, 2019 | 2019 | http://hdl.handle.net/11250/2582811 | SDS011 | Y | Y | http://inovafitness.com/en/a/chanpinzhongxin/95.html | OPC | OEM |  | PM2.5 | µg/m3 |  | black box | FIELD | traffic |  | 1 | 1 | hour | [7 empty] | RF | [9 empty] | TEOM | Sensor | linear | 0.76 | 2.21 | 0.775 | [28 empty] | 0.4 - 127.5 | ug.m-3 | [15 empty] | 30 | EUR |  |  | JRC-FK
Figaro |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | MOs | OEM | Figaro TGS-5042 | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.27 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.022 | -11.08 | 57.6 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
Figaro |  |  | Spinelle et at. 2017 Part B | 2017 | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf | TGS-5042 | Y | Y |  | MOs | OEM | Figaro TGS-5042 | CO | ppm | Y | open source | FIELD | semi-rural | < 60 |  | 1 | hour | 0.3 | [6 empty] | MLR | Rs=a1*CO+a2*T+a3*RH+a0 |  | T | RH | [5 empty] | REF | Sensor | linear | 0.036 | 8.15 | -39 | [45 empty] | 40 | EUR | http://db-airmontech.jrc.ec.europa.eu/download/1-s2.0-S092540051631070X-main.pdf |  | JRC-FK
TZOA |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | TZOA-PM Research Sensors | Y | Y |  | nephelometer | OEM |  | PM | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.4356 | 1.37 | 6.68 | 17 |  |  |  | 47 | [40 empty] | 90 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
TZOA |  |  | Feinberg et al. 2018 | 2018 | https://www.atmos-meas-tech.net/11/4605/2018/ | TZOA-PM Research Sensors | Y | Y |  | nephelometer | OEM |  | PM | hppcf | Y | open source | FIELD | urban |  |  | 1 | hour | [17 empty] | REF | Sensor | linear | 0.518 | 2.16 | 6.75 |  |  |  |  | 47 | [40 empty] | 90 | USD | https://www.atmos-meas-tech.net/11/4605/2018/ |  | JRC-FK
UB AirSense-Sharp |  |  | Manikonda et al. 2016 | 2016 | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | UBAS | Y | N |  | nephelometer | sensor system | Sharp-GP2Y1010AU0F | PM2.5 | µg/m3 | N | black box | LAB |  |  |  | 1 | min | [17 empty] | REF | Sensor | linear | 0.85 | 40 | 1.6 |  |  |  | 4 | [41 empty] | 100 | USD | https://linkinghub.elsevier.com/retrieve/pii/S0021850216301021 | cigarette smoke | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5038 | 13.514 | 0.9981 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.4827 | 14.702 | 1.1936 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | GRIMM | linear | 0.5333 | 11.741 | 1.2428 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2497 | 29.319 | 0.5408 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2254 | 30.739 | 0.6207 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | GRIMM | linear | 0.2558 | 27.366 | 0.7088 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2769 | 28.249 | 0.8719 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2756 | 28.708 | 1.0624 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.2882 | 26.985 | 1.0764 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6821 | 14.313 | 1.1983 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.6716 | 16.516 | 1.4124 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.7045 | 13.314 | 1.4761 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4542 | 22.042 | 0.7617 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4502 | 22.757 | 0.9234 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 5 | min | [17 empty] | Sensor | T640 | linear | 0.4713 | 21.075 | 0.9467 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4615 | 22.376 | 0.8111 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4514 | 23.045 | 0.98 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 1 | hour | [17 empty] | Sensor | T640 | linear | 0.4858 | 210.23 | 1.0069 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7066 | 11.438 | 1.0602 | [6 empty] | 38 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7076 | 13.102 | 1.2603 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
Alphasense |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 | OPC-N3 | Y | Y | https://airthinx.io/products/iaq/ | OPC | OEM |  | PM10 | µg/m3 | N | black box | FIELD | rural |  | 10 | 24 | hour | [17 empty] | Sensor | T640 | linear | 0.7444 | 10.193 | 1.3189 | [6 empty] | 25 | [38 empty] | 338 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/alphasense-opc-n3---field-evaluation.pdf?sfvrsn=6 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7593 | 3.5122 | 0.5884 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7493 | 4.4044 | 0.5785 |  |  |  |  | 99.5 |  | 11 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.7585 | 4.1737 | 0.5652 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8621 | 3.4227 | 0.5845 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8479 | 4.3302 | 0.5717 |  |  |  |  | 99.5 |  | 11 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM2.5 | µg/m3 | N | black box | FIELD | rural |  | 35 | 24 | hour | [17 empty] | Sensor | BAM | linear | 0.8595 | 4.1342 | 0.5562 |  |  |  |  | 99.5 |  | 12 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4563 | 9.4619 | 17.916 |  |  |  |  | 99.5 |  | 1.2 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.362 | 12.243 | 31.138 |  |  |  |  | 99.5 |  | 0.7 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
PurpleAir |  |  | AQ-SPEC | 2018 | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 | PA-I-Indoor | Y | Y | https://www.wunderground.com/sensors/purpleair | OPC | sensor system | Plantower-PMS5003 | PM10 | µg/m3 | N | black box | FIELD | rural |  | 35 | 1 | hour | [17 empty] | Sensor | BAM | linear | 0.4608 | 10.151 | 14.728 |  |  |  |  | 99.5 |  | 1.4 | [38 empty] | 180 | USD | http://www.aqmd.gov/docs/default-source/aq-spec/field-evaluations/purple-air-pa-i-indoor---field-evaluation.pdf?sfvrsn=13 |  | JRC-FK
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | min | [17 empty] | TEOM | Sensor | linear | 0.48 |  |  |  |  | 73 | [17 empty] | 54.5 | [24 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 15 | min | [17 empty] | TEOM | Sensor | linear | 0.525 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 1 | hour | [17 empty] | TEOM | Sensor | linear | 0.565 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
Winsen |  |  | Badura et al. 2018 | 2018 | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html | ZH03A | Y | Y |  | OPC | OEM |  | PM2.5 | µg/m3 | Y | black box | FIELD | urban |  |  | 24 | hour | [17 empty] | TEOM | Sensor | linear | 0.67 |  |  |  |  | 73 | [42 empty] | 20 | USD | https://www.e3s-conferences.org/articles/e3sconf/abs/2018/19/e3sconf_eko-dok2018_00006/e3sconf_eko-dok2018_00006.html |  | JRC-FK
HabitatMap | IMT Lille Douai | AirCasting | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirBeam | Y | Y | http://aircasting.org/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 1 | 16.4009 | -0.2901 | 2.4804 |  |  |  | 0.1984 | [12 empty] | 0 | SD | [26 empty] | 200 | USD |  |  | INERIS-LSp
HabitatMap | IMT Lille Douai | AirCasting | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirBeam | Y | Y | http://aircasting.org/ | nephelometer | sensor system | Shineyi PPD60PV | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5394 | 14.9891 | -0.873 | 0.7207 |  |  |  | 0.6944 | [12 empty] | 1.6608 | SD | [26 empty] | 200 | USD |  |  | INERIS-LSp
ENVEA Cairpol | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/fr/envea/ | electrochemical | sensor system | Membrapor | NO2 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0076 | 355.9855 | 2.3694 | 0.5389 |  |  |  | 76.9841 | [12 empty] | 189.6988 | SD | [26 empty] | 820 | EUR |  |  | INERIS-LSp
ENVEA Cairpol | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/fr/envea/ | electrochemical | sensor system | Membrapor | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8658 | -3.3197 | 0.7878 | 0.0067 |  |  |  | 99.4048 | [12 empty] | 2.3973 | SD | [26 empty] | 820 | EUR |  |  | INERIS-LSp
ENVEA Cairpol | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | CairClip NO2-F | Y | Y | http://cairpol.com/fr/envea/ | electrochemical | sensor system | Membrapor | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.4678 | -1.8726 | 0.2202 | 0.0153 |  |  |  | 81.5476 | [12 empty] | 1.6475 | SD | [26 empty] | 820 | EUR |  |  | INERIS-LSp
Alphasense | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3125 | -1.4886 | 2.8756 | 0.0573 |  |  |  | 84.5238 | [12 empty] | 41.5542 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0158 | -11.3267 | 7.4656 | 0.8703 |  |  |  | 84.5238 | [12 empty] | 729.3728 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | IMT Lille Douai |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0908 | 7.1269 | 2.6047 | 0.0986 |  |  |  | 84.5238 | [12 empty] | 78.1606 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Normandie Tech'Air |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirMatrix | Y | Y | https://fr-fr.facebook.com/AirmatrixFrance/ | nephelometer | sensor system |  | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.9511 | -0.2928 | 1.0292 | 0.0146 |  |  |  | 12.6984 | [12 empty] | 1.8555 | SD | [26 empty] | 60 | EUR |  |  | INERIS-LSp
Normandie Tech'Air |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirMatrix | Y | Y | https://fr-fr.facebook.com/AirmatrixFrance/ | nephelometer | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8149 | -5.7724 | 0.9292 | 0.0579 |  |  |  | 14.4841 | [12 empty] | 4.9923 | SD | [26 empty] | 60 | EUR |  |  | INERIS-LSp
Normandie Tech'Air |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirMatrix | Y | Y | https://fr-fr.facebook.com/AirmatrixFrance/ | nephelometer | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.84 | -1.542 | 1.0914 | 0.0294 |  |  |  | 14.4841 | [12 empty] | 4.151 | SD | [26 empty] | 60 | EUR |  |  | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7013 | -1.1543 | 0.8988 | 0.0052 |  |  |  | 99.6032 | [12 empty] | 4.54 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0982 | 11.8592 | 0.303 | 0.0138 |  |  |  | 99.6032 | [12 empty] | 10.2272 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.632 | -0.9874 | 1.9186 | 0.0295 |  |  |  | 99.9008 | [12 empty] | 18.3292 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0133 | 10.2088 | 1.7629 | 0.3604 |  |  |  | 99.9008 | [12 empty] | 242.5666 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.4088 | 0.1593 | 1.9993 | 0.0518 |  |  |  | 99.9008 | [12 empty] | 30.3785 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3515 | 3.0086 | 0.4255 | 0.0076 |  |  |  | 99.9008 | [12 empty] | 4.4675 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7435 | 3.1773 | 0.5917 | 0.0079 |  |  |  | 99.9008 | [12 empty] | 3.8649 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6413 | -0.8685 | 2.0467 | 0.0306 |  |  |  | 99.9008 | [12 empty] | 19.1645 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2958 | 2.598 | 1.5505 | 0.0526 |  |  |  | 99.9008 | [12 empty] | 38.2182 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.4788 | -0.3515 | 2.1319 | 0.0432 |  |  |  | 99.9008 | [12 empty] | 28.1036 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6552 | 1.6736 | 1.0309 | 0.0084 |  |  |  | 68.75 | [12 empty] | 5.4138 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2425 | 4.8884 | 0.5871 | 0.0179 |  |  |  | 68.75 | [12 empty] | 10.6954 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2554 | 3.0051 | 1.0871 | 0.0242 |  |  |  | 68.254 | [12 empty] | 16.1978 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | http://www.ecomesure.com/fr/systemes-connectes/ecomsmart | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0159 | 5.0182 | 1.2883 | 0.1607 |  |  |  | 68.254 | [12 empty] | 110.1475 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
ECOMESURE |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ECOMSMART | Y | Y | https://eshop.ecomesure.com/index.php?id_product=1&id_product_attribute=0&rewrite=ecomsmart&controller=product&id_lang=2 | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.1232 | 6.2195 | 0.9829 | 0.0352 |  |  |  | 68.254 | [12 empty] | 23.4883 | SD | [26 empty] | 4650 | EUR |  | Price 3400 + 300 (NO2, O3) + 800 (PM) + 200 (WIFI) + 250 (Battery) + 100 (Wall mounting) | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6607 | 4.469 | 1.1864 | 0.0162 |  |  |  | 60.3175 | [12 empty] | 6.7511 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2803 | 3.4765 | 0.8444 | 0.0286 |  |  |  | 47.4206 | [12 empty] | 14.677 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.023 | -4.3302 | 1.9618 | 0.2872 |  |  |  | 61.5079 | [12 empty] | 178.3659 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0004 | 2.6602 | 0.007 | 0.0231 |  |  |  | 61.6071 | [12 empty] | 3.6416 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6436 | 0.762 | 1.0009 | 0.0074 |  |  |  | 77.6786 | [12 empty] | 5.8322 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5229 | 0.1104 | 1.0122 | 0.018 |  |  |  | 63.8889 | [12 empty] | 10.7096 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0137 | -2.8833 | 1.5166 | 0.2327 |  |  |  | 77.9762 | [12 empty] | 182.2669 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0037 | 1.8989 | -0.0144 | 0.0211 |  |  |  | 78.0754 | [12 empty] | 2.5108 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.609 | 0.6712 | 1.0437 | 0.0071 |  |  |  | 82.9365 | [12 empty] | 5.8136 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5726 | -2.6455 | 1.0976 | 0.0159 |  |  |  | 69.0476 | [12 empty] | 10.1544 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0061 | -2.092 | 1.3462 | 0.2131 |  |  |  | 82.8373 | [12 empty] | 216.8027 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0054 | 3.8337 | -0.0302 | 0.0166 |  |  |  | 83.3333 | [12 empty] | 3.9817 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | http://www.libelium.com/products/plug-sense/technical-overview/ | MOs | sensor system |  | NO | ppb | Y | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7705 | -1.6554 | 0.9837 | 0.0136 |  |  |  | 79.6627 | [12 empty] | 10.7658 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | https://www.iot-store.com.au/products/libelium-waspmote-plug-sense-smart-environment-pro | MOs | sensor system | Waspmote Plug & Sense - Smart Environment PRO | NO2 | ppb | Y | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6687 | 9.2829 | 1.7395 | 0.0225 |  |  |  | 96.0317 | [12 empty] | 9.4909 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | http://www.libelium.com/products/plug-sense/technical-overview/ | MOs | sensor system |  | NO2 | ppb | Y | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.47 | 105.7111 | 2.3211 | 0.1286 |  |  |  | 96.0317 | [12 empty] | 19.1074 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | http://www.libelium.com/products/plug-sense/technical-overview/ | MOs | sensor system |  | NO2 | ppb | Y | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6536 | 48.1923 | 1.9096 | 0.0632 |  |  |  | 96.0317 | [12 empty] | 10.7759 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | http://www.libelium.com/products/plug-sense/technical-overview/ | OPC | sensor system | Alphasense | PM1 | ppb | N | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2998 | 0.2664 | 0.8772 | 0.0162 |  |  |  | 96.0317 | [12 empty] | 15.5855 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | http://www.libelium.com/products/plug-sense/technical-overview/ | OPC | sensor system | Alphasense | PM10 | ppb | N | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0377 | 6.7075 | 2.487 | 0.1981 |  |  |  | 96.0317 | [12 empty] | 186.7604 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Libelium | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Waspmote | Y | Y | https://www.cooking-hacks.com/shop/sensors/gas | OPC | sensor system | Alphasense | PM2.5 | ppb | N | open source | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0958 | 7.8209 | 1.4066 | 0.0536 |  |  |  | 96.0317 | [12 empty] | 49.846 | SD | [26 empty] | 1270 | EUR |  | Price 490 + 80 (4 MOs) + 700 (alphaSense OPC, factor 2), check with Laurent | INERIS-LSp
Rubix Sense & Instrumentation | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.9138 | 0.5562 | 1.5003 | 0.009 |  |  |  | 96.3294 | [12 empty] | 3.5758 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8497 | 3.1574 | 1.4315 | 0.0107 |  |  |  | 96.3294 | [12 empty] | 4.6708 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5019 | -9.4821 | 3.2275 | 0.0498 |  |  |  | 96.3294 | [12 empty] | 36.8912 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0317 | -24.1346 | 5.9995 | 0.5121 |  |  |  | 96.3294 | [12 empty] | 484.8788 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3776 | -11.3107 | 3.7602 | 0.0711 |  |  |  | 96.3294 | [12 empty] | 54.8303 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7708 | 2.2072 | 1.2439 | 0.0079 |  |  |  | 96.4286 | [12 empty] | 5.2608 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5911 | 1.6888 | 1.0339 | 0.0124 |  |  |  | 82.1429 | [12 empty] | 9.5922 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0106 | 1.2753 | 0.4784 | 0.0711 |  |  |  | 96.4286 | [12 empty] | 67.9212 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0079 | 2.5036 | -0.0229 | 0.0175 |  |  |  | 96.4286 | [12 empty] | 2.9394 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2131 | 11.0643 | 1.1479 | 0.0223 |  |  |  | 95.9325 | [12 empty] | 17.108 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | electrochemical | sensor system |  | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2359 | 10.4818 | 1.1973 | 0.0365 |  |  |  | 84.127 | [12 empty] | 24.0393 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0102 | 1.5359 | 0.6531 | 0.0976 |  |  |  | 96.4286 | [12 empty] | 94.401 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Vaisala | Atmo Grand Est |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQT-420 | Y | Y | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/aqt420 | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.011 | 1.9427 | -0.022 | 0.0178 |  |  |  | 96.4286 | [12 empty] | 2.3802 | SD | [26 empty] | 5000 | EUR |  |  | INERIS-LSp
Rubix Sense & Instrumentation |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7225 | 4.9489 | 1.2362 | 0.0283 |  |  |  | 33.2341 | [12 empty] | 5.4914 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | electrochemical | sensor system | Alphasense | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0423 | 45.9436 | -0.153 | 0.0673 |  |  |  | 33.2341 | [12 empty] | 6.9145 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5672 | -2.7082 | 2.7676 | 0.0541 |  |  |  | 33.2341 | [12 empty] | 12.609 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0411 | -154.4183 | 19.5334 | 2.2364 |  |  |  | 33.2341 | [12 empty] | 735.2928 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
Rubix Sense & Instrumentation |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Watchtower 1 | Y | Y | https://www.rubixsi.com/fr/environnement-exterieur/watchtower-1/ | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.1551 | -7.0516 | 3.2537 | 0.1229 |  |  |  | 33.2341 | [12 empty] | 37.9365 | SD | [26 empty] | 5000 | EUR |  | There is also the cost of software for managing several units (4000 Euro) | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM1 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.9072 | -0.5403 | 1.1266 | 0.0046 |  |  |  | 83.4325 | [12 empty] | 3.5427 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM10 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5608 | 1.8135 | 1.2443 | 0.0185 |  |  |  | 83.4325 | [12 empty] | 13.9088 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6196 | 2.7456 | 1.5129 | 0.0184 |  |  |  | 83.4325 | [12 empty] | 11.5154 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM1 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.9095 | -0.528 | 1.1531 | 0.0048 |  |  |  | 83.4325 | [12 empty] | 3.5768 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM10 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5619 | 1.595 | 1.5362 | 0.0257 |  |  |  | 83.4325 | [12 empty] | 17.1345 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6015 | 2.929 | 1.546 | 0.0196 |  |  |  | 83.4325 | [12 empty] | 12.2215 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM1 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.9807 | -0.6915 | 1.3143 | 0.0149 |  |  |  | 33.4325 | [12 empty] | 1.8828 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM10 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3441 | 16.6346 | 0.8756 | 0.0605 |  |  |  | 33.4325 | [12 empty] | 14.7015 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
AtmoTrack by 42 Factory |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Atmotrack | Y | Y | https://atmotrack.fr/en/index.html | OPC | sensor system | Chinese manufacturer | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 30 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3725 | 17.8289 | 0.9403 | 0.0636 |  |  |  | 33.4325 | [12 empty] | 13.1407 | SD | [26 empty] | 2500 | EUR |  | CO, NH3 et NO2, MEMS + electrochimique in Version 2 | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7307 | -4.9906 | 1.8514 | 0.0141 |  |  |  | 89.881 | [12 empty] | 8.8818 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6581 | 4.2695 | 1.818 | 0.0326 |  |  |  | 85.7143 | [12 empty] | 14.6138 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.4632 | 0.5975 | 0.602 | 0.0128 |  |  |  | 83.2341 | [12 empty] | 8.2776 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3327 | -0.1428 | 1.1782 | 0.0278 |  |  |  | 99.9008 | [12 empty] | 27.5373 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7729 | -0.7016 | 0.9299 | 0.0084 |  |  |  | 83.2341 | [12 empty] | 6.744 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3971 | 30.256 | 3.246 | 0.0774 |  |  |  | 83.5317 | [12 empty] | 27.8578 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7084 | 24.0303 | 3.9514 | 0.1228 |  |  |  | 82.8373 | [12 empty] | 27.1952 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8214 | -0.2922 | 0.2884 | 0.0159 |  |  |  | 83.2341 | [12 empty] | 1.7179 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0482 | 5.2398 | 1.6485 | 0.1221 |  |  |  | 99.9008 | [12 empty] | 120.8488 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3147 | 1.0601 | 0.7796 | 0.015 |  |  |  | 99.9008 | [12 empty] | 14.6842 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | NO2 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3419 | 17.7216 | 1.2236 | 0.0318 |  |  |  | 74.0079 | [12 empty] | 12.0969 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | electrochemical | sensor system | Alphasense | O3 | ppb | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.6865 | 6.4459 | 1.2708 | 0.022 |  |  |  | 74.504 | [12 empty] | 9.285 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.7884 | -0.3751 | 0.4494 | 0.0172 |  |  |  | 66.5675 | [12 empty] | 3.1347 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.06 | 3.3474 | 1.6336 | 0.1079 |  |  |  | 99.9008 | [12 empty] | 106.684 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Environmental Instruments Ltd | ADDAIR |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AQMesh v.4.0 | Y | Y | https://www.aqmesh.com/ | OPC | sensor system | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3449 | 1.0951 | 0.9683 | 0.0169 |  |  |  | 99.9008 | [12 empty] | 17.0334 | SD | [26 empty] | 8000 | EUR |  |  | INERIS-LSp
Met One | Envicontrol |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Met One (NM) | Y | Y | https://metone.com/products/npm-2-network-particulate-monitor/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 3600 | 1 | hour | [17 empty] | Sensor | Reference | linear | 0.8466 | -2.6592 | 1.1028 | 0.0075 |  |  |  | 68.254 | [12 empty] | 4.8742 | SD | [26 empty] | 1900 | USD |  |  | INERIS-LSp
Met One | Envicontrol |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | Met One (NM) | Y | Y | https://metone.com/products/npm-2-network-particulate-monitor/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 3600 | 1 | hour | [17 empty] | Sensor | Reference | linear | 0.8646 | -2.821 | 1.1283 | 0.0072 |  |  |  | 68.254 | [12 empty] | 4.6353 | SD | [26 empty] | 1900 | USD |  |  | INERIS-LSp
Met One | Envicontrol |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ES-642 | Y | Y | https://metone.com/products/es-642/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8294 | -0.0023 | 0.0011 | 0.4042 |  |  |  | 5.2579 | [12 empty] | 0.0073 | SD | [26 empty] | 2600 | USD |  |  | INERIS-LSp
Met One | Envicontrol |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | ES-642 | Y | Y | https://metone.com/products/es-642/ | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.8442 | -0.0023 | 0.0011 | 0.3773 |  |  |  | 6.3492 | [12 empty] | 0.0069 | SD | [26 empty] | 2600 | USD |  |  | INERIS-LSp
liberaintentio | Atmo Normandie |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/airsenseur-air-quality-monitoring-open-framework/ | electrochemical | sensor system | Membrapor NO/C-25 | NO |  | Y | black box | FIELD | urban |  | 1 | 1 | min | [7 empty] | linear | [9 empty] | Sensor | Reference | linear | 0.0321 | 8744.3479 | 1.4106 | 145.7565 |  |  |  | 16.5675 | [12 empty] | 75.1826 | SD | [26 empty] | 1600 | EUR |  |  | INERIS-LSp
liberaintentio | Atmo Normandie |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/airsenseur-air-quality-monitoring-open-framework/ | electrochemical | sensor system | AlphaSense NO2-B43F | NO2 |  | Y | black box | FIELD | urban |  | 1 | 1 | min | [7 empty] | linear | [9 empty] | Sensor | Reference | linear | 0.7307 | 62205.1726 | -7.1069 | 61.7305 |  |  |  | 99.9008 | [12 empty] | 33.3555 | SD | [26 empty] | 1600 | EUR |  |  | INERIS-LSp
liberaintentio | Atmo Normandie |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | AirSensEUR (v.1) | Y | Y | https://airsenseur.org/website/airsenseur-air-quality-monitoring-open-framework/ | electrochemical | sensor system | AlphaSense OX-A431 | O3 |  | Y | black box | FIELD | urban |  | 1 | 1 | min | [7 empty] | linear | [9 empty] | Sensor | Reference | linear | 0.2275 | 62008.2517 | -1.671 | 80.3611 |  |  |  | 83.2341 | [12 empty] | 35.4202 | SD | [26 empty] | 1600 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.3269 | 2.3537 | 1.7322 | 0.0573 |  |  |  | 37.004 | [12 empty] | 19.1004 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.1365 | 8.7391 | 1.1037 | 0.084 |  |  |  | 37.004 | [12 empty] | 29.6111 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.2224 | 3.8281 | 1.6362 | 0.0744 |  |  |  | 37.004 | [12 empty] | 25.2993 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM1 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.1861 | 4.1976 | 0.9065 | 0.0297 |  |  |  | 81.8452 | [12 empty] | 24.2931 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM10 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0083 | 16.7166 | 0.6641 | 0.1455 |  |  |  | 81.8452 | [12 empty] | 119.5958 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Alphasense | Atmo Auvergne-Rhône-Alpes |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | OPC-N2 | Y | N | http://www.alphasense.com/index.php/products/optical-particle-counter/ | OPC | OEM | Alphasense | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 1 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0681 | 8.4558 | 0.8532 | 0.0503 |  |  |  | 81.8452 | [12 empty] | 41.0228 | SD | [26 empty] | 362 | EUR |  |  | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.5105 | 1.4932 | 0.5951 | 0.0786 |  |  |  | 23.5119 | [12 empty] | 12.0764 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.662 | 0.9343 | 0.6992 | 0.0485 |  |  |  | 23.5119 | [12 empty] | 8.5249 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.141 | 10.9222 | 0.229 | 0.1004 |  |  |  | 23.7103 | [12 empty] | 11.6713 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0919 | 10.6514 | 0.1751 | 0.0775 |  |  |  | 23.7103 | [12 empty] | 9.3778 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.108 | 14.0267 | 0.2697 | 0.0995 |  |  |  | 23.8095 | [12 empty] | 16.0076 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0648 | 13.7996 | 0.1964 | 0.0796 |  |  |  | 23.8095 | [12 empty] | 12.727 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM10 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0606 | 23.7012 | 0.2785 | 0.1128 |  |  |  | 23.8095 | [12 empty] | 22.6582 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Groupe TERA |  |  | LCSQA DRC_18_174307_09689A | 2018 | https://www.lcsqa.org/system/files/rapport/LCSQA2017-CILmicrocapteurs-synthese_resultats.pdf | e-PM | Y | Y | http://groupe-tera.com/pdf/fr/e-PM-AQSense.pdf | OPC | sensor system |  | PM2.5 | µg/m3 | N | black box | FIELD | urban |  | 60 | 1 | min | [17 empty] | Sensor | Reference | linear | 0.0201 | 21.6398 | 0.1428 | 0.0939 |  |  |  | 23.8095 | [12 empty] | 17.0093 | SD | [26 empty] | 2500 | EUR |  | Suivant la durée et le nombre de capteurs, le prix peut varier de 375€/semaine à 385€/mois (soit environ 95 €/semaine) | INERIS-LSp
Intel Labs Berkeley |  | https://www.communitysensing.org | Vaughn D et al. 2010 (Sonoma Technology) | 2010 | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf | Intel Berkeley Badge | Y | N |  | MOs | sensor system | SGX Sensortech - MICS-2610 | O3 | ADC Units | Y | open source | LAB |  |  |  | 1 | hour | 0.8794 | 521.76 | 368.75 |  |  |  |  | log | Y = a0 + a1 ln(x) | [8 empty] | REF | Sensor | logaritmic | [49 empty] | EUR | https://db-airmontech.jrc.ec.europa.eu/download/NO2%20Sensors%20Report.pdf |  | JRC-FK

## references
reference
kumar_rise_2015
cen_ambient_2012_CO
cen_ambient_2012_NO2
cen_ambient_2012_O3
cen_ambient_2012_SO2
cen_ambient_2014_PM
lewis_validate_2016
spinelle_evaluation_2017
key-vocs_metrology_2017
Spinelle_et_al_2013
Spinelle_et_al_2013(a)
lewis_low-cost_2018
barrett_coefficient_1974
the_world_air_quality_index_sensing_2019
iscape_summary_2017
aleixandre_review_2012
snyder_changing_2013
white_sensors_2012
williams_air_2014
directive_2008_50_ec
williams_deliberating_2019
fishbain_evaluation_2017
esposito_dynamic_2016
european_commission_guide_2010
thunis_performance_2012
de_vito_calibrating_2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Borghi et al. 2018
Borghi et al. 2018
Borghi et al. 2018
Borghi et al. 2018
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Jiao et al. 2016
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
US-EPA_2015
Williams et al. 2014(b)
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Zikova et al. 2017
Zikova et al. 2017
Zikova et al. 2017
Zikova et al. 2017
Zikova et al. 2017
Zikova et al. 2017
Williams et al. 2014(c)
Williams et al. 2014(c)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
Sun et al. 2016
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Mijling et al. 2017
Vaughn D et al. 2010 (Sonoma Technology)
Vaughn D et al. 2010 (Sonoma Technology)
Vaughn D et al. 2010 (Sonoma Technology)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Wastine_2019_a
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et al. 2015 (a)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Zimmerman_et_al_2018
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Mueller et al. 2017
Bigi et al. 2018
Spinelle et at. 2017 Part B
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Wastine_2019_a
Bigi et al. 2018
Spinelle et at. 2017 Part B
Spinelle et at. 2015 Part A (b)
Marjovi et al. 2017
Marjovi et al. 2017
Marjovi et al. 2017
Cordero et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Bigi et al. 2018
Cordero et al. 2018
Bigi et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
Cordero et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Karagulian et al. 2019
Wastine_2019_a
Wastine_2019_a
Cordero et al. 2018
Wastine_2019_a
Wastine_2019_a
Cordero et al. 2018
Wastine_2019_a
Wastine_2019_a
Cordero et al. 2018
Wastine_2019_a
Wastine_2019_a
Cordero et al. 2018
Wastine_2019_b
Wastine_2019_b
Wastine_2019_a
Wastine_2019_b
Wastine_2019_b
Wang et al. 2015
Alvarado et al. 2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Mead et al. 2013
Mead et al. 2013
Mead et al. 2013
Mead et al. 2013
Mead et al. 2013
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
Popoola et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Cavaliere et at. 2018
Cavaliere et at. 2018
Cavaliere et at. 2018
Cavaliere et at. 2018
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Castell et all. 2017
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Cross et al. 2017
Cross et al. 2017
Cross et al. 2017
Cross et al. 2017
Bettair 2017
Bettair 2017
Bettair 2017
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Spinelle et al. 2015 (a)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Duvall at al. 2016
Spinelle et al. 2015 (a)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Williams et al. 2014(c)
Duvall at al. 2016
Duvall at al. 2016
Feinberg et al. 2018
Feinberg et al. 2018
Williams et al. 2014(b)
Williams et al. 2014(b)
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Williams et al. 2014(b)
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Gerboles et al. 2015(a)
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Chakrabarti et al. 2004
Chakrabarti et al. 2004
Chakrabarti et al. 2004
AQ-SPEC
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
Viana et al. 2015
Sousan et al 2016
Williams et al. 2014(b)
Williams et al. 2014(b)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Sousan et al 2016
Northcross et al. 2013
Northcross et al. 2013
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Steinle et al. 2015
Steinle et al. 2015
Han et al,.2016
Han et al,.2016
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Jovasevic_Stojanovic et al 2015
Wastine_2019_b
Northcross et al. 2013
Northcross et al. 2013
Northcross et al. 2013
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
US-EPA_2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
Gillooly et al. 2019
Gillooly et al. 2019
Gillooly et al. 2019
Gillooly et al. 2019
Spinelle et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Olivares and Edwards, 2015
Manikonda et al. 2016
Manikonda et al. 2016
Manikonda et al. 2016
Sousan et al 2016
Alvarado et al. 2015
Alvarado et al. 2015
Wang et al. 2015
Borrego et al. 2016
Borrego et al. 2016
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
Kunak field report 2017
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Piedrahita et al. 2014
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Sun et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Williams et al. 2014(b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Williams et al. 2014(c)
Williams et al. 2014(c)
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Borrego et al. 2016
Borrego et al. 2016
Borrego et al. 2016
Spinelle et at. 2017 Part B
Zimmerman_et_al_2018
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Gerboles et al. 2015(a)
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Spinelle et al. 2015 (a)
Gerboles et al. 2015(a)
Spinelle et al. 2015 (a)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Sun et al. 2017
Gerboles et al. 2015(a)
Spinelle et al. 2015 (a)
Spinelle et at. 2015 Part A (b)
Spinelle et at. 2015 Part A (b)
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Wei et al. 2018
Spinelle et at. 2015 Part A (b)
Wastine_2019_a
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Dacunto et al. 2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Mukherjee_et_al_2017
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Sousan et al. 2016 (a)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Di-Antonio_developing_2018
Di-Antonio_developing_2018
Di-Antonio_developing_2018
Di-Antonio_developing_2018
Di-Antonio_developing_2018
Di-Antonio_developing_2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Crilley et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Pillarisetti et al 2017
Pillarisetti et al 2017
Pillarisetti et al 2017
Pillarisetti et al 2017
Williams et al. 2014(c)
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Williams et al. 2014(b)
Williams et al. 2014(b)
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
Kelly et al 2017
Kelly et al 2017
AQ-SPEC
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Zheng et al. 2018
Kelly et al 2017
Laquai_particle_2017
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Wang et al. 2015
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
Holstius et al 2014
AQ-SPEC
Gao et al. 2015
Gao et al. 2015
Kelly et al 2017
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Zimmerman_et_al_2018
Zimmerman_et_al_2018
AQ-SPEC
Zimmerman_et_al_2018
Zimmerman_et_al_2018
Kelly et al 2017
Lin et al. 2015
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Vaughn D et al. 2010 (Sonoma Technology)
Budde et al. 2018
Budde et al. 2018
Budde et al. 2018
Austin et al. 2015
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Liu et al, 2019
Laquai_particle_2017
Liu et al, 2019
Liu et al, 2019
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Karagulian et al. 2012
Feinberg et al. 2018
Feinberg et al. 2018
Feinberg et al. 2018
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Liu et al, 2019
Spinelle et at. 2017 Part B
Spinelle et at. 2017 Part B
Feinberg et al. 2018
Feinberg et al. 2018
Manikonda et al. 2016
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
AQ-SPEC
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
Badura et al. 2018
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
LCSQA DRC_18_174307_09689A
Vaughn D et al. 2010 (Sonoma Technology)

## Sheet2
Honeywell HPM | HPM series | optical | sensor | PM2.5 | N | black box |  |  | up to 95% |  |  | 15ug/m3 (o-100 ug/m3); 15% 100-1000 ug/m3 | 0-1000 ug/m3 |  |  | https://sensing.honeywell.com/sensors/particle-sensors | 22 | EUR
Winsen ZH03A | ZH03 - Laser Dust Module | optical | sensor | PM2.5 | N | black box |  |  |  | PM2.5=1000x(TH)/(TH+TL) (low & high lev. Time) | [5 empty] | https://www.winsen-sensor.com/d/files/air-quality/zh03-series-laser-dust-module-v2_0.pdf | 72 | EUR
PLANTOWER PMS A003 | [15 empty] | http://www.plantower.com/content/?125.html
PLANTOWER PMS 1003
Shinyei | PM Sensor Evaluation Kit | optical | sensor system | PM2.5 | N | black box | 0.92 | [8 empty] | https://www.shinyei.co.jp/stc/eng/optical/main_pmmonitor.html | 5200 | EUR
Dylos | DC 1700 | optical | sensor system | PM2.5 | N | black box | 0.917 | [8 empty] | http://www.dylosproducts.com/dc1700.html | 425 | EUR
PMX | Custom PMX | optical | sensor system | PM2.5 | N | open source | 0.97 | 0.1905 |  |  | 30.8 ug/m3 (std) |  |  |  |  | Park_et_al_2013
Air nut sensor |  | optical | sensor system | PM2.5 | N | black box | 0.81 | [6 empty] | 20-23 | 20-23 | file:///C:/Users/karagfe/Downloads/SDS011_laser_PM2.5_sensor_specification-V1.3.pdf | 200 | EUR
Nova Fitness Co. Ltd. | SDS011 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m3 |  | 0-1000 ug/m3 |  |  | https://pan.baidu.com/s/1skCk8AL | 20 | EUR
Nova Fitness Co. Ltd. | SDS018 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m4 |  | 0-1000 ug/m4 |  |  | https://ecksteinimg.de/Datasheet/SDS018%20Laser%20PM2.5%20Product%20Spec%20V1.5.pdf | 20 | EUR
Nova Fitness Co. Ltd. | SDS021 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m5 |  | 0-1000 ug/m5 |  |  | https://cdn.sparkfun.com/assets/parts/1/2/2/7/5/SDS021_laser_PM2.5_sensor_specification-V1.0.pdf | 20 | EUR
Nova Fitness Co. Ltd. | SDS019 | optical | sensor | PM2.5 | N | open source |  |  | 70% |  | 15% and ± 15ug/m6 |  | 0-1000 ug/m6 |  |  | https://pan.baidu.com/s/1sl7Hj0h | 21 | EUR
Met One | ES-642 | optical | sensor | PM2.5 | N | open source | 0.92
Winsen | ZE03-NO2 | electrochemical | sensor | NO2 | N | open source | [9 empty] | https://aqicn.org/air/view/sensor/spec/no2.winsen-ZE03.pdf | 90 | EUR
Winsen | ZE03-CO | electrochemical | sensor | CO | N | open source | [9 empty] | https://aqicn.org/air/view/sensor/spec/no2.winsen-ZE03.pdf | 90 | EUR
Alphasense | NO2-A1
https://www.epa.gov/air-sensor-toolbox/evaluation-emerging-air-pollution-sensor-performance
EPA reports
https://www.epa.gov/air-sensor-toolbox/air-sensor-toolbox-resources-and-funding#TER
https://www.epa.gov/sites/production/files/2018-03/documents/final_em-3_master_slide_set.pdf

## Notes
Columns
NOTES
1 Sensor system is intented to be a sensor together with a DAQ (data acquisition) and data transfer; Sensor is intented to be the measuring device only.
3Parameter used for the MLR in addtion to the reference
2Rs is refered to the signal response of the sensor, NO2, O3, NO, CO, T, RH are the intereferents
OEM = Original Equipment Manufacturer
| Response time | t90, time necessary to reach 90 % of the final value
| Resolution time | Lowest time for acqusition of a new data
Bias
| Range | tange of values of the measured pollutant: minimum - maximum
| MOdel type for calibration | type of calibration model equation applied to raw data of the sensor, when empty raw data are use, MLR: multi linear regression, RF, random forest, ANN: artifical neural networ, log: logarithmic, exp: exponential
BL | U | masurement uncertainty or standard deviation
BM | type of U | "U at concentration level" or "SD"