# Review of performance of Low-cost Sensors for Air Quality Monitoring

Federico Karagulian, Maurizio Barbiere, Alexander Kotsev, Laurent Spinelle , Michel Gerboles1, Friedrich Lagler, Nathalie Redon, Sabine Crunaire and Annette Borowiak

last update 05 July 2019, 16:47

## Abstract

A growing number of companies started commercializing Sensor Systems able to give a “best guess” for the concentration of indoor and outdoor pollutants. The benefit about the use of low-cost sensor is the increase spatial coverage when monitoring air quality in cities and remote locations. Today, there are more than 250 low-cost sensor systems commercially available on the market with a cost ranging from a few hundreds to a few thousand euro. At the same time, independent information about the performance of sensor systems against reference measurements is only available for about 70 sensor systems in literature. In fact, the behaviour of commercial low-cost sensors is unstable and often affected by atmospheric condition, pollutants concentration range and, by the site location where the measurements are carried out. In this work, quantitative data about the performance of tested low-cost sensors against reference measurement were collected into a repository. This information was gathered from published reports and relevant testing laboratories. Other information was drawn from peer-reviewed journals that tested different types of sensors in research studies. Relevant metrics about the comparison of sensor systems against reference systems highlited the the most cost-effective sensor systems that could be used to monitor air quality pollutants with a good level of agreement rapresented by a coefficient of determination of R2 > 0.75 a slope of 1.0 ± 0.5 and a price < 3 k€. This review wanted to highlight the possibltiy to have versatile sensors able to operate with multiple pollutants and with transparent data treatment.

All references here1–86

## 1. Introduction

The introduction and diffusion of the micro-sensors technology for monitoring ambient air pollution (as emerging measuring devices) is contributing to the rapid adoption of low-cost sensors (LCS) for air quality monitoring for citizen science initiative and by public authorities40. Low-cost in this context is typically referring to the cost of the hardware component needed to make a measurement. LCS can provide real time measurements at lower cost allowing higher spatial coverage than the current reference methods of measurements of air pollutants. Additionally, the monitoring of air pollution with reference measurements methods requires skilled operators for the maintenance and calibration of measuring devices that are described in detailed Standard Operational Procedures14–18. Conversely, it should be possible that LCS are operated without human intervention making it possible for unskilled users to be able to monitor air pollution without the need of important technical understanding.

However, a lot of LCS are becoming available whose performance regarding the agreement between LSC values and reference measurement can be of variable quality making it fundamental to evaluate LCS before choosing any LCS for routine measurements and case studies44. As reported below, few independent tests are reported in academic publications. The rapid technological progress of LCS and the time needed to publish studies in academic journals makes publication of articles not the preferred route and the majority of the available information is found in grey literature, mainly of report types.

The major sources of information of LCS evaluation consist of AQ-SPEC3, the US-EPA and the work carried out by the Joint Research Centre62. (check in the article US-EPA workshop what do they say). Although a number of reviews of the suitability of sensors for ambient air quality have been published1,12,33,40,58,79,80,84, quantitative data for comparing and evaluating the agreement between sensors and reference data are mostly missing. Additionally, there is no commonly accepted protocol for the test of LCS, the metrics reported are generally diverse making it difficult to compare the performance of sensor between evaluation studies. The most common reported metrics consist of: the coefficient of determination, R2, the slope and intercept of regression line between sensor and reference measurements, the Root mean square of Error, RMSE and the measurement uncertainty . Here after, the results of an exhaustive review of the existing literature on LCS evaluation are presented.

The purpose of this review is to identify LCSs whose comparison with reference measurements shows the highest correlation and accuracy. For this purpose, we performed a comprehensive review about the performance of commercial LCSs. We have aggregated summary statistics about the agreement between sensors and reference instruments. Although in Europe, the main metrics to evaluate the performance of measuring methods consists of the measurement uncertainty, this metrics could not be used in our study since the majority of studies do not report it (287 records out of 1423 total number of records reporting RMSE and other metrics for uncertainty ) give the number of studies with this parameter?). Conversely, we had to rely on most common metrics, i. e., the coefficient of determination R2, the slope and intercept of linear regression line between sensor and reference measurement and, in few cases, the Root Mean Square of Error that were scrutinized and analyzed to identify sensors that could potentially be complementary to the reference methods of air quality monitoring.

The market of LCSs for ambient air monitoring only consists of a small number of sensor model types that are manufactured by a few companies. These LCSs, are usually known as Original Equipment Manufacturers (OEM).

As per definition, an OEM is a chemical cell or physical unit that produces an analytically useful signal by detecting or measuring the analyte. On the other hand a Sensor System (SS) or sensor node is an integrated set of hardware that uses one or more sensors to detect and/or measure a chemical concentration or quantity that is able to supply real time measurements. A sensor systems contain a number of common components in addition to the basic sensing/analytical element that is used for detection. Common core components and functions may include:

- Sensing element or detector (actually the sensor)

- Sampling capability (active or passive sampling)

- Power systems, including batteries

- Analogue to digital conversation

- Signal processing

- Local data storage

- Data transmission

- Remote calibration

- Housing/casing

OEMs use chemical and physical techniques phenomena to sense pollutant in ambient air. However, in order to simplify measurement operations, calibration and data transfer into a convenient sensor object, OEMs need be integrated into a sensor system (SS), consisting of electronic boards, software and protective box gathering the hardware, software and OEM sensors.

The use of low-cost sensors is extensively interesting for citizen-science initiatives. Therefore, Small Medium Enterprises were able to sell sensor-systems which could be deployed by citizen who wanted to monitor air quality in a chosen environment. Up to date, there are several sensor systems using sensors from the same OEM. However, outputs from these sensors system often differ from each other. The ideal candidate sensor system would show good agreement with reference measurements and, at the same time, provide sensor raw data allowing to be calibrated using open source correction algorithms. The number of air pollutants being measured was also a parameter taken into consideration. Finally, the price of a low cost-sensor was also taken into account.

## 2. Methods

### 2.1. Data sources

About 1423 records were systematically gathered from peer-reviewed studies of sensors for air quality and air pollution reported in the Scopus database, the World Wide Web, the AirMontech web site (http://db-airmontech.jrc.ec.europa.eu/search.aspx), ResearchGate and Google search. Overall, about a number of 64 independent studies were found from different sources from reports, peer-review papers and sensors manufacturers. The research was focused on sensors for Particulate Matter (PM), Ozone (O3), Nitric Dioxide (NO2) and Carbon Monoxide (CO). A few references were also included for nitrogen monoxide sensors (NO). We have started evaluating summary statistics from the correlation of sensors with reference measurements (validation).

The research covered the period between 2010 and 2018 (year of publication). Data gathered from sensor studies were reviewed according to criteria described in the following sections. Additionally, reviewed sensor data were used to populate a database that was used to generate summary statistics about characteristics and performances of sensors.

### 2.2. Data collection

Most of the reviewed studies reported regression parameters obtained from the comparison between sensors and reference measurements. Records from regression parameters were the result of both calibration and comparison of the sensor with a reference instrument. In the case of calibration, we identified four types of most used regression models: linear, multilinear (MLR), quadratic and logarithmic.

### 2.3. Evaluation criteria

We have carried out an extensive literature review of OEM sensors and sensor systems (SS) that were used to estimate concentration of air pollutants against a reference systems during field and laboratory tests. The purpose was to gather quantitative information about the performance of sensors according to the following criteria:

- Agreement between sensor and reference measurements

- Availability of raw data, transparency of data treatment and possibility of a-posteriori calibration

- Capability to measure multiple pollutants

- Affordability of sensor systems taking into consideration the number of provided sensors

- Capacity to satisfy the requirement of interoperability of data according to the INSPIRE directive

- Automatic data-transfer and website visualization of sensor data

The review was focused on available commercially available OEMs and sensor systems, even though a few non-commercial sensor were also considered, measuring concentrations of Particulate Matter (PM2.5, PM10, PM1), Nitric Dioxide (NO2), Nitrix Monoxide (NO), Carbon Monoxide (CO) and Ozone O3. Table 1 reports the number of records, by pollutant and operating technology, gathered in literature about validation and testing of OEMs / sensor systems against a reference system. Records were collected from laboratory (133) and field tests (1290).

Table 1. Number of analyzed records for OEMs/Sensor Systems by pollutant and by type of technology.

| pollutant | type | field___lab | n. records | references |
| --- | --- | --- | --- | --- |
| CO | electrochemical | FIELD | 51 | AQ-SPEC[3], Jiao[33], Sun[69], Wastine[76], Zimmerman[85], Marjovi[48], Karagulian[36], Popoola[57], Castell[12], Borrego[10], Cross[22], Gillooly[29] |
| CO | electrochemical | LAB | 9 | Sun[69], Mead[49], Castell[12], Gerboles[28], Wei[78], Zimmerman[85] |
| CO | MOs | FIELD | 27 | AQ-SPEC[3], Spinelle[64], Borrego[10], Piedrahita[55] |
| CO | MOs | LAB | 2 | AQ-SPEC[3], Piedrahita[55] |
| NO | electrochemical | FIELD | 44 | AQ-SPEC[3], Jiao[33], Bigi[8], Wastine[76], Spinelle[64], Karagulian[36], Mead[49], Popoola[57], Castell[12], Borrego[10], Cross[22], Gillooly[29], LCSQA[42] |
| NO | electrochemical | LAB | 6 | Castell[12], Gerboles[28], Wei[78] |
| NO | MOs | FIELD | 1 | LCSQA[42] |
| NO2 | electrochemical | FIELD | 137 | AQ-SPEC[3], Jiao[33], Sun[69], Mijling[50], Spinelle[63], Mueller[51], Bigi[8], Marjovi[48], Cordero[20], Karagulian[36], Wastine[76], Wastine[77], Mead[49], Popoola[57], Borrego[10], Castell[12], Cross[22], Duvall[25], Gillooly[29], Zimmerman[85], LCSQA[42] |
| NO2 | electrochemical | LAB | 21 | Williams[81], Sun[69], Vaughn[73], Castell[12], Spinelle[61], Gerboles[28], Wei[78], Sun[70], Zimmerman[85] |
| NO2 | MOs | FIELD | 28 | AQ-SPEC[3], US-EPA[72], Borrego[10], Piedrahita[55], Spinelle[63], Lin[45], LCSQA[42] |
| NO2 | MOs | LAB | 10 | Vaughn[73], Williams[81], Piedrahita[55] |
| O3 | electrochemical | FIELD | 65 | Jiao[33], Spinelle[63], Mueller[51], Karagulian[36], Wastine[76], AQ-SPEC[3], Borrego[10], Castell[12], Cross[22], Duvall[25], Feinberg[26], LCSQA[42] |
| O3 | electrochemical | LAB | 10 | Spinelle[61], Castell[12], Gerboles[28], Wei[78] |
| O3 | MOs | FIELD | 54 | AQ-SPEC[3], Jiao[33], Marjovi[48], Borrego[10], Feinberg[26] |
| O3 | MOs | LAB | 3 | AQ-SPEC[3], Spinelle[62], Vaughn[73] |
| O3 | UV | FIELD | 9 | AQ-SPEC[3] |
| O3 | UV | LAB | 1 | Sun[69] |
| PM2.5 | Electrical | FIELD | 6 | AQ-SPEC[3] |
| PM2.5 | nephelometer | FIELD | 129 | Borghi[9], Jiao[33], Feinberg[26], US-EPA[72], Williams[80], AQ-SPEC[3], Zikova[84], Chakrabarti[19], Borrego[10], Olivares[54], Holstius[31], Gao[27], Karagulian[35], LCSQA[42] |
| PM2.5 | nephelometer | LAB | 24 | Manikonda[47], AQ-SPEC[3], Wang[75], Alvarado[2], Sousan[59], Holstius[31], Kelly[37], Austin[4] |
| PM2.5 | OPC | FIELD | 428 | AQ-SPEC[3], Mukherjee[52], Feinberg[26], Jiao[33], Cavaliere[13], Williams[80], Borrego[10], Viana[74], Northcross[53], Holstius[31], Steinle[68], Han[30], Jovasevic[34], Gillooly[29], Sun[69], Dacunto[23], Crilley[21], Di-Antonio[24], Badura[5], Pillarisetti[56], Kelly[37], Zheng[83], Laquai[41], Budde[11], Liu[46], LCSQA[42] |
| PM2.5 | OPC | LAB | 27 | AQ-SPEC[3], Cavaliere[13], Manikonda[47], Northcross[53], Sousan[60], Pillarisetti[56], Kelly[37] |
| PM1 | Electrical | FIELD | 6 | AQ-SPEC[3] |
| PM1 | nephelometer | FIELD | 1 | LCSQA[42] |
| PM1 | OPC | FIELD | 102 | AQ-SPEC[3], Williams[80], Crilley[21], Di-Antonio[24], LCSQA[42] |
| PM1 | OPC | LAB | 8 | AQ-SPEC[3], Sousan[60] |
| PM10 | nephelometer | FIELD | 26 | AQ-SPEC[3], Borrego[10], LCSQA[42] |
| PM10 | nephelometer | LAB | 1 | Alvarado[2] |
| PM10 | OPC | FIELD | 176 | AQ-SPEC[3], Cavaliere[13], Borrego[10], Feinberg[26], Han[30], Jovasevic[34], Williams[80], Crilley[21], Budde[11], LCSQA[42] |
| PM10 | OPC | LAB | 11 | AQ-SPEC[3], Cavaliere[13], Manikonda[47], Sousan[59], Sousan[60] |

Most of laboratory and field tests were collected results published and between the year 2010 and 2018. As reported in Figure 1, only few preliminary works were published from 2010 to 2014. On 2015 we recorded the highest number of references with 16 different works publishing results about performances of sensors for air quality monitoring. Since 2015, the number of references publishing works on sensors with about 10 - 12 publications per year.

Figure 1. Number of references per year of publication.

Overall, we found 34 references reporting field tests with sensors co-located at urban sites, 7 references for rural sites and, 10 references for traffic sites. Most of the laboratory and field tests reported hourly averages with about 610 records obtained for over 86 models of OEMs and sensor systems. On the other hand, we found about 248 records from tests performed over an averaging time of 24 hours and 5 minutes with about 42 models of sensors (Table S1). Therefore, 1 hour averaged measurements were considered statistically more signifative to represent most of the scrutinized 112 models of OEMs and sensor systems.

For the evaluation of sensors against reference systems we considered the metrics that were more frequently reported in the reviewed works about the evaluation of the sensor performance both in field and/or laboratory tests. The coefficient of determination R2 is usually used as indication of “usefulness” or “goodness” of fit obained from regression models comparing sensor measurements with reference measurements Table S2. However, R2 is a partial measures of how sensors data agree with reference measurements according to a regression model6. A larger R2 reflects an increase in the predictive precision of the regression model but it ignores the loss in information due to possible loss in degrees of freedom. A significance test is therefore suggested in this case. Alternatively, R2 can be viewd as a measure of goddness of fit (how close evaluation data is to the reference measurements) and the slope of the regression as level of accuracy. However, if the goodness of fit about the regression is fixed, then the slope will increase and consequentely also the R2. Therefore, when it happens to calibrate different datasets, calibration using slope and R2 values close to 1.0 might be less precise than calibration using smaller values of slope and R2. As shown in Table 2, most of the works reported coefficients of determination R2 as well as the slope of regression line mainly from field tests. Root mean square error (RMSE) and the uncertainty (U), were mostly calculated as value of standard deviation and only reported in few works (28). Therefore, for the purpose of this work, we only focused on the analysis of the comparions of R2 and slope of laboratory and field tests of sensors.

Table 2. Number of records gathered by metric used in this work.

| metrics | n. Field Tests | n. Laboratory Tests |
| --- | --- | --- |
|  | 1290 | 133 |
| R2 from calibrations | 218 | 60 |
| R2 from comparisons | 1164 | 72 |
| slope of reg. line | 1063 | 55 |
| intercept | 1027 | 54 |
| RMSE | 285 | 5 |
| Uncertainity (U) | 153 | 29 |

## 3. Classification of sensors

For each model of sensor system we identified the manufacturer of OEM sensor and the manufacturer of the sensor system itself including the sensors, protective box, internal hardware and software for data acquisition, data treatment and data transfer. Each database record describing the laboratory or field performance of a sensors was considered valuable only if it reported information about laboratory tests and/or field comparison against a reference measurement at a reference monitoring station. Overall, we found 112 models of sensors including both OEMs (31) and sensor systems (81) manufactured by 77 manufacturers (16 OEM and 61 SS). In addition, we identified 19 projects about the evaluation of OEMs/sensor systems under different operational conditions (Air Quality Egg, Air Quality Station, AirCasting,3,9,27,52, Carnegie Mellon27,85, CitiSense82 Cairsense34, Developer Kit3, HKEPD/14-0277169, making-sense.eu50, communitysensing.org73, MacPoll.eu63, OpenSense II8,51, Proof of Concept AirSensEUR36, SNAQ Heathrow49,57). Out of 1423 records collected from literature, we identified 1192 records (201 OEM and 991 SS) from 90 alive sensors (25 OEM and 65 SS) and 231 records (119 OEM and 112 SS ) from 22 “non active” (or discontinued) sensors (6 OEM and 16 SS).

Commonly speaking, “low-cost” refers to the cost of a sensor system compared to the cost of a reference instrument measuring the air pollutant43. More recently, ultra-affordable OEMs are starting to appear on the market for PM monitoring. Many of them are designed to be integrated in the Internet of things (IoT) network of devices. Currently, for the detection of PM2.5, it is possible to purchase optical sensor at prices of few hundreds euros to few tens of euros from devices manufactured in emerging economies such as the Republic of China and Republic of Korea71. For the detection of PM2.5, some of these sensors are starting achieving performances comparable to low-cost OEMs manufactured in the Western world.

### 3.1. Raw data and traceability of data treatment

The data treatment of measurements of OEM sensors and sensor systems can be classified in two distinct categories:

- Data acquisition and data processing of sensor data that are operated by an “open source” software tuned according to different calibration parameters and environmental conditions. All data treatments from data acquisition until the conversion to pollutant concentration levels is known to the user. Among 34 sensors, we identified 234 records from OEMs (108 ) and sensor systems (126) using an open source software for data management. Usually, outputs from these sensor were already in the same measurement units as the reference measurements.In this category, sensor devices are generally connected to a custom-made data acquisition system to acquire sensor raw data. Generally, users are expected to set a calibration function in order to convert sensor raw data to validate against reference measurements. In this category, the sensor only consists of OEMs integrated in data acquisition system, in most cases with a posteriori data treatment.

- Sensor systems with integrated built-in OEM sensors, data acquisition system and calibration algorithms whose data treatment is unknown and without the possibility to change any parameter have been identified as “black box”. This is because of the impossibility for the user to accurately know the whole chain of data treatment and to tuen the sensor itself. Among 83 sensors, we identified 1189 records among OEMs (212) and sensor systems (977) not using an open source software for data management.In most cases, the sensor system wasare previously calibrated against a reference system or, the calibration parameters wereare remotely adjusted by the manufacturer.

Clear definitions and examples of the principle of operations used by the different type of sensor (electrochemical, metal oxides, optical particulate counter, optical sensors) are reported in a recent work by WMO43. This work also reports the several limitation of each type of sensor such as, interference by meteorological parameters, cross-sensitivities to other pollutants, drifts and aging effect. At the date, there is a larger number of active and commercially available sensors (Figure 2). However, while most of the OEM sensors are open sources, allowing end-users to integrate them into sensors systemes, most of the sensor systems themselves were found to be “black-box” devices. This represents a limitation when the sensor system might need further re-calibration other then the one provided by the manufacturer.

Sensor are also classified according to their commercial availability. Sensor were assigned to the “Commercial” category if they could be purchased and operated by any user without the need for another service for data treatment or reporting activing (as a consultant service for example). Sensors fell under category non-commercial when it was non possible to find any supplier for purchasing. Typically this type of sensor are used for research and publication while it is difficult for any user to repeat the same sensor setup.

Figure 2. Number of sensor models gathered from the literature review. Sensors has ben classified by their type of technology, availability, openness and commerciality.

### 3.2. Sensors for air quality

For the detection of Particulate Matter, the largest number of sensor tests, carried out with OEM and SS, was found for Optical Particle Counters (OPC) with 752 records followed by Nephelometers with 181 records (Table 1). Both systems detect particulate matter by measuring the light scattered by particles, with the OPC being able to directly count particles according to their size. On the other hand, nephelometers estimate particle density that is subsequently converted into particle mass. For the detection of gaseous pollutant such as, NO2, CO and O3, the largest number of tests was found for electrochemical sensors with 343 records, followed by metal oxides sensors (MOs) with 125 records (Table 2). Electrochemical sensors are based on a chemical reactions between gases in the air and the working electrode of an electrochemical cells. The current of electrons being created is measured between the working and counter electrode that are dipped into an electrolyte in a metal oxide sensor (also named resistive sensor, semiconductor). Therefore, gases in the air react on the surface of a semi-conductor and exchange electrons modifying its conductance.

Table S3 reports the models of OEM sensors currently used to monitor particulate matter according to principle of operation and the cut-off of particulate size. For the sake of clarity, we aggregated records from PM2.5 together with PM0.5−2.5, PM2 and PM3. The same was done for PM10 records that were aggregated with coarse PM2.5−10. On the other hand, models of sensor systems measuring concentration of particulate matter were reported in Table S42–5,9–11,13,19,21,23,24,27,28,30–32,34–38,41,42,47,52–54,56,59,68,74,75,81–83,85. Several sensor systems can use the same OEM. In very few cases, the same model of sensor system was tested using different types of OEM sensors when performing validation tests. Table S5 and Table S6 report the models of OEM and sensor systems, respectively, currently used to measure concentration of gaseous air pollutants NO2, O3, NO and, CO according their their type of technology.3,7,8,10,12,13,20,22,23,26,27,29,30,34,41,48–51,55,61,63,64,66,67,69,70,72,73,78,82,86

## 4. Calibration of Sensors

Calibration of the sensor is somewhat considered a sensitive information from most of sensor manufacturers. Several studies performed calibration of sensors during laboratory or field tests. The calibration consisted in the application of a regression model in order to adjust the response of the sensor to a reference system. We found calibration records for both OEMs and sensor systems. Overall, we gathered over 352 records about calibration of sensors using different types of mathematical models (Table 3) and at different time resolutions. The linear model and the multi linear regression model (MLR) were largely used to calibrate the sensor response against a reference measurement. Other calibration approaches used the exponential, logarithmic, quadratic, Random Forest and, few types of neural network models. We could observe that most of MLR models, covariates such as meteorological parameters Temperature and Relative Humidity, and gaseous pollutant such as, Nitric Dioxide (NO2), Nitric Monoxide (NO) and Ozone (O3), were used to optimize the calibration. Some types of model also took into cosnideration the time−drift of the sensor as covariate. The calibration of OEMs was performed using the raw signal of the sensor that most of the time was expressed as a voltage or as a current. On the other hand, for sensor systems, the calibration was carried out using the units of the reference system.

Table 3. Types of calibration models used for the calibration of sensors at different time resolutions (ANN: artificial neural network, exp: exponential; log: logarithmic; MLR: multilinear regression; quad: quadratic; RF: random forest; SVM: support vector machine; SVR: support vector regression)

| pollutant | calibration model | n. records | references | median_r2_calib | median_r2 |
| --- | --- | --- | --- | --- | --- |
| CO | ANN | 2 | Wastine[76], Spinelle[64] | NA | 0.58 |
| CO | linear | 12 | Sun[69], Wastine[76], Castell[12], Cross[22], Gerboles[28], Spinelle[64], Zimmerman[85] | 0.85 | 0.15 |
| CO | MLR | 21 | Jiao[33], Karagulian[36], Wastine[76], Wei[78], Piedrahita[55], Spinelle[64], Zimmerman[85] | 0.89 | 0.83 |
| CO | quad | 12 | AQ-SPEC[3] | 0.63 | NA |
| CO | RF | 1 | Zimmerman[85] | 0.91 | NA |
| NO | ANN | 2 | Wastine[76], Spinelle[64] | NA | 0.57 |
| NO | linear | 8 | Wastine[76], Castell[12], Cross[22], Spinelle[64], Gerboles[28], LCSQA[42] | 0.96 | 0.032 |
| NO | MLR | 20 | Jiao[33], Bigi[8], Karagulian[36], Wastine[76], Spinelle[64], Wei[78] | 0.92 | 0.91 |
| NO | RF | 2 | Bigi[8] | NA | 0.9 |
| NO | SVR | 2 | Bigi[8] | NA | 0.90 |
| NO2 | ANN | 7 | Spinelle[63], Cordero[20], Wastine[76], Wastine[77] | 0.87 | 0.94 |
| NO2 | linear | 25 | Sun[69], Vaughn[73], Spinelle[63], Wastine[76], Wastine[77], Castell[12], Cross[22], Gerboles[28], Zimmerman[85], Lin[45], LCSQA[42] | 0.29 | 0.17 |
| NO2 | log | 1 | Vaughn[73] | 0.89 | NA |
| NO2 | MLR | 48 | Jiao[33], Sun[69], Mijling[50], Spinelle[63], Mueller[51], Bigi[8], Cordero[20], Karagulian[36], Wastine[76], Wastine[77], Piedrahita[55], Wei[78], Sun[70], Zimmerman[85] | 0.81 | 0.81 |
| NO2 | quad | 6 | AQ-SPEC[3] | 0.61 | NA |
| NO2 | RF | 7 | Bigi[8], Cordero[20], Zimmerman[85] | 0.86 | 0.91 |
| NO2 | SVM | 4 | Cordero[20] | 0.85 | 0.94 |
| NO2 | SVR | 2 | Bigi[8] | NA | 0.78 |
| O3 | ANN | 2 | Spinelle[63], Wastine[76] | NA | 0.89 |
| O3 | linear | 13 | AQ-SPEC[3], Sun[69], Spinelle[63], Wastine[76], Castell[12], Cross[22], Gerboles[28], LCSQA[42] | 0.84 | 0.53 |
| O3 | log | 1 | Vaughn[73] | 0.88 | NA |
| O3 | MLR | 20 | Jiao[33], Spinelle[63], Karagulian[36], Wastine[76], Spinelle[62], Wei[78] | 0.91 | 0.88 |
| O3 | quad | 9 | AQ-SPEC[3] | 0.72 | NA |
| PM1 | Kholer | 2 | Di-Antonio[24] | NA | 0.74 |
| PM1 | log | 6 | AQ-SPEC[3] | 0.76 | NA |
| PM10 | exp | 6 | AQ-SPEC[3] | 0.59 | NA |
| PM10 | Kholer | 2 | Crilley[21] | NA | NA |
| PM10 | linear | 3 | AQ-SPEC[3], Cavaliere[13], Jovasevic[34] | 0.77 | 0.63 |
| PM10 | log | 7 | AQ-SPEC[3] | 0.58 | NA |
| PM10 | quad | 1 | Alvarado[2] | 0.65 | NA |
| PM10-2.5 | linear | 4 | Sousan[59], Han[30], Jovasevic[34] | 0.63 | 0.98 |
| PM2.5 | exp | 3 | Dacunto[23], Kelly[37], Austin[4] | 0.91 | 0.97 |
| PM2.5 | Kholer | 4 | Crilley[21], Di-Antonio[24] | NA | 0.78 |
| PM2.5 | linear | 36 | Mukherjee[52], Wang[75], Alvarado[2], Cavaliere[13], Jovasevic[34], Olivares[54], Kelly[37], Zheng[83], Holstius[31] | 0.84 | 0.67 |
| PM2.5 | log | 7 | AQ-SPEC[3], Laquai[41] | 0.73 | NA |
| PM2.5 | MLR | 17 | Jiao[33], Sun[69], Zheng[83], Holstius[31], Liu[46] | 0.81 | 0.65 |
| PM2.5 | quad | 8 | Chakrabarti[19], Alvarado[2], Zheng[83], Gao[27] | 0.87 | 0.88 |
| PM2.5 | RF | 3 | Liu[46] | NA | 0.79 |
| PM2.5-0.5 | linear | 9 | Northcross[53], Steinle[68], Han[30], Jovasevic[34] | 0.84 | 0.98 |
| PM2.5-0.5 | MLR | 1 | Jiao[33] | 0.6 | 0.45 |
| PM2.5-0.5 | quad | 6 | AQ-SPEC[3], Manikonda[47] | 0.82 | NA |

As explained above, from the analyzed records, we found several type of regression model that were used to calibrate sensors from OEM and sensor systems against reference systems. In order the estimate quality of the used calibration model, we reported the coefficient of determination R2 as indication of the amount of total variability explained by the model. On a first instance, The coefficient of determination can be used as indication of performance of the calibration model chosen to validate the sensor with a reference system. In addition to simple linear models, raw sensor data were validated using multilinear and quadratic models which included the use of covariates to improve the quality of the calibration (Table 3).

Figure 3 shows a summary of all mean R2 obtained from the calibration of sensor systems against reference measurements. Results were grouped by model of sensor system and averaged per reference work. For the same sensor systems we can observe R2 ranging from low to high values up to the unit. This show the variability of the performance of sensor systems depending from the type of calibration. In the following we are going to report a detailed discussion about performance of several OEM and sensor systmens during their calibration at different averaging-times.

Figure 3. Mean R2 for obtained from the calibration of sensor systems against reference measurements.

Calibration of sensor data against a reference system was found to be carried out using input data at different time resolution. Therefore, in order to make a comparison of R2 obtained at the same time resolution data, we chose records averaged over different time-scale of 1 hour (Figure S1) and 1 minute (Figure S2). Most of these records were from OEMs (85) whereas only a limited number were from sensor systems (109). For the measurement of PM2.5, values of R2 ~ 1 were found for the sensors PMS1003 by Plantower38 at 1-hour resolution and for the the PMS3003 , Dylos DC1100 PRO and DC1700 by Dylos at a resolution of 1 minute.3,68,83 The Plantower and Dylos sensors showed higher R2 when calibrated with 1 minute resolution reference data. Other sensors such as, the OPC-N2 by AlphaSense3 reported values of R2 falling within the range of 0.7 - 1.0 at a resolution of 1 hour. The same OEM sensor OPC-N2, reported values of R2 just above 0.7 when measuing PM1 while it did not show a good performance when measuring PM103. We need to stress out that optical sensors, such as OPCs and nephelometers, are somewhat limited when detecting coarse particulate patter PM10 because of the low-efficiency of the sampling system when sampling large particles in ambient air.

Most of regression models used for the calibration of sensors detecting gaseous pollutants used a time-resolution of 1 hour. For the calibration of sensors measuring O3, the largest values of R2 was reported for the OEM sensors FIS SP-61 by FIS and O3-3E1F by CityTechnology, when using a time-resolution of 1 hour (Figure S1)61. On the other hand, when using a time-resolution of 1 minute, values of R2 ~ 1 were found for the sensor system AirSensEUR (v.2) by LiberaIntentio36 as well as for the OEM S-500 by Aeroqual3 (Figure S2). The AirSensEUR uses a built-in OEM OX-A431. We want to point out that, most of the MLR models used for calibration ozone sensors foresees the use of reference NO2 because of the strong oxidizing effect of O3 on gas sensors with consequent formation of NO2. For the calibration of sensors measuring NO2 we found values of R2 within the range 0.7 - 1.0 for the OEM sensor NO2-B42F (by Alphasense78), at a time resolution of 1 hour, and the sensor systems AirSensEUR (v.2) (by LiberaIntentio)36) and MAS69 at a time resolution of 1 minute (Figure 3). We need to point out that for the measurement of NO2, the AirSensEUR (v.2) uses the OEM sensor NO2-B43F by AlphaSense.

Most of the records about the calibration of sensor measuring CO showed high values of R2. As shown in Figure S1, the OEMs CO 3E300 by City Technology29 and CO-B4 by Alphasense78 reported R2 ~ 1 for time-resolution of 1 hour. High values of R2 were also reported for the sensor system AirSensEUR (v.2) when calibrating for CO at a time-resolution of 1 minute (Figure S2)36. Other sensors reporting values of R2 within the range 0.7 - 1.0 where the MICS-4515 by and SGX Sensortech55, the Smart Citizen Kit by Acrobotic3 and the RAMP86. All these sensors used 1 hour time-resolution data.

## 5. Comparison with reference systems

We found about 1236 records about the comparison of calibrated sensors against a reference instrument. All comparisons were carried out by using a linear regression model between calibrated and reference data. The performance of the regression was evaluated with the coefficient of determination R2 and the slope of the regression. As explained above, we need to stress out that not all the analyzed records reported the Root Mean Square Error (RMSE) of the regression therefore, we decided to omit it in the present review.

In this work, records gathered from the comparison of sensors with reference systems came from OEMs and sensor systems using a custom calibration or a built-in calibration directly setup by the manufacturer. As for the records collected from the calibration of sensor, comparison with reference system was carried out at different time-resolutions. Here we only report comparisons performed at a time-resolution of 1 hour with 565 and 151 records from sensor systems and OEMs, respectively. In Figure 4 we have reported the averaged R2 for all reviewd sensor systems found from each reference workd. As for the evaluation of the calibration of sensors, we have averaged R2 valuse obtained from the same raference and for the same sensor systems. It was observed that, for the same sensor system, we could find different R2 value for the comparison with reference measurements. In the following we give a more detalied breakdown about the performances of each OEM / SS upon their comparison with reference mesurements at different averaging-times.

Figure S3 shows the distribution of R2 for sensors systems measuring PM2.5, PM10, PM1, O3, NO2 and CO against reference at 1-hour time-resolution. For the measurements of particulate matter, most of the comparisons were performed during field tests with the highest R2 obtained from the sensor PA-II by PurpleAir3 and PATS+ by Belkley Air56. These sensors reported values of R2 between 0.8 and 1.0. Other sensors with R2 values falling in the range 0.7-1.0 were identified in the PMS-SYS-1 by Shinyei, the Dylos 1100 PRO by Dylos, the MicroPEM by RTI, the AirNUT by Moji China  the Egg (2018) by Air Quality Egg, the AQT410 v.1.15 by Vaisala, the AirVeraCity by AirVeraCity, the NPM2 by MetOne23 and, the Air Quality Station by AS LUNG3. Records from other sensors showed different values of R2 depending of the type of field test and for the averaging time chosen to process the time-series of data. We need to point out that the performance of sensor systems measuring PM10, on average, was very poor.

For gaseous pollutants, high R2 values were found for the sensor systems 2B Tech. (POM) by 2B Technologies (O3)3, the AirSensEUR (v.2) by LiberaIntentio36 the Bettair by Bettair Cities7 the AirCasting by HabitatMap81 the KUNAKAIR A10 V2 by kunak41 (O3, NO2, CO and NO), the Spec, the AQMesh. These sensors reported values of R2 between 0.8 and 1.0. As shown in Figure S3, we found a non-negligible number of records for sensor systems whose R2 resulting from the comparison with reference systems was within the range 0.7 - 1.0 using 1-hour averaged data. We want to point out that, among all tested sensor systems, only the AirSensEUR (v.2) was the only one measuring multiple pollutants.

The comparison of OEMs against reference systems, showed only few sensors for the measurement of PM had R2 within the range 0.7 - 1.0 when average over a time-scale of 1 hour. Among them we could identify the Dylos DC 170032,35,47,59 and the OPC-N2 (Figure S4)3,5,21,27,52 when measuring PM2.5. On the other hand, when the comparison was performed over a time-scale of 24 hour we found R2 within the range 0.7 - 1.0 for several OEMs which included the PMS7003 by Plantower5, the SDS011 by Nova Fitness5, the OPC-NO23, and the Egg v.2 (PM) by Air Quality Egg3 when measuring PM2.5 (Figure S5). The same behavior was observed from the comparison of sensor systems against a reference system when measuring 24-hour averaged data of PM2.5. As we can see from Figure S6, several sensor systems such as, Dylos DC 170053, PA-II,3 AirQUINO by CNR13 reported values of R2 ~ 1.

Figure 4. Mean R2 for obtained from the comparison of sensor systems against reference measurements.

For the evaluation of gaseous pollutants, we found very few OEMs with R2 within 0.7 - 1.0 when using data at at time-resolution of 1 hour. These sensors included the CairClip O3/NO2 by CairPol26,27,63,82, the Aeroqual Series 500 (and SM50)27, the O3-3E1F by CityTechnology27,29,63,65 and the NO2-B43F by Alphasense70,86 (Figure S4). On the other hand, we found very few records for sensor systems using 24 hour data. As a general remark, we can see that the performance of OEMs sensors is enhanced when they are integrated inside a sensor systems. It is also evident that most of the gathered records from PM2.5 and gaseous pollutants O3, NO2, CO and NO, used 24 hours and 1 hour time-resolution data as required by the European Air Quality Directive.

To check the accuracy of a sensor, when compared to a reference system, we looked at the value of the slope obtained from the linear regression of the sensor measurements against a reference measurement. Most of comparisons were carried out during field tests, while only a limited number laboratory tests were available. Ideally, only an R2 ~ 1.0 and a slope ~ 1.0 should be a good indicator of performance for a sensor. Therefore, we only selected records with R2 > 0.7 and slope within the range 0.5-1.5.

Figure S7 shows sensor systems such as, the AQM 60, the KUNAKAIR A10 V2, the AirSensEUR (v2) has slopes ~ 1 for most of measured gaseous pollutants when using 1 hour time-resolution data. On the other hand, only few records from PM2.5 sensor systems showed slopes ~ 1 for 1-hour (PATS+ and AirNut) and 24-hour (AIRQuino) time-averaged sensor systems. Only few PM2.5 sensors performed well at ideal conditions. Sensor systems such as, AIRQino, SidePak AM510, Air Quality Egg (v.2) (PM), Dylos DC1100 PRO, AirNut and the OEM OPC-N2 were in good agreement with a reference system (Figure S7, Figure S8).

Among OMEs showing slopes ~ 1 when using 1-hour time-averaged data, we found the SM50, the CairClip O3/NO2, the S-500, (NO2, O3) and, the NO2-B4F (NO2) and the Egg v.2 (PM) (PM2.5) (Figure S9). On the other hand, when using 24-hour time-averaged data, the OEM OPC-N2 by Alphasense and the Egg v.2 (PM) by Air Quality Egg, shoed slopes ~ 1 when measuring PM10 (Figure S10).

As general remark, from the above analysis we could observe that for some OEMs and sensor systems, the width of the interquartile range IQR (H-spread) was very narrow. This is an indication of the reproducibility of the regression parameters used in their calibration. This becomes relevant when it comes to the development of a reliable sensor system that uses the same OEM sensor and the same calibration algorithms. From the present analysis,

## 6. Price of sensors

As preliminary outcome of the present analysis, we have started identifying sensors systems that are in good agreement with reference instruments commonly used to monitor pollutant concentrations. Although the sensor market constantly develops, we decided to identify a sensor system that is commercially available or that can be assembled with commercially available OEM sensors.

Usually, the price of OEM sensors only represents a small fraction of the selling value of the entire sensor system. In the common understanding, a sensor for air quality is classified as low-cost when its price is less than 10000 EUR. In addition, if a low-cost sensor can measure multiple pollutants, potentially it could be used by local authorities as complementary source of air quality data as subsitute of reference instruments whose cost might rise up to one order of magnitude.

For the evaluation of the price of sensors, we considered all sensor systems manufactured by commercial companies as well as sensor systems built for laboratory testing by research groups. The latter ones are custom-built devices assembled around an OEM sensor. We must to stress out that, while for the detection of different size of particulate matter it is possible to use the same optical sensor, for the detection of gaseous pollutant it is necessary to have a dedicated sensor for each pollutant. Therefore, among all the analyzed records, we tried identifying sensor systems that can measure concentration of particulate matter together with gaseous pollutants.

Figure S12 and Figure 5 show the commercial price of OEMs and sensor systems by model and number of pollutant measured by each sensor. There is a large number of sensor systems measuring single pollutants but few ones measuring multiple pollutants. This is an indication about the complexity to have a sensor system measuring multiple pollutants. Most of OEMs are open source devices (Figure S12). This means that OEMs can be used to build sensor systems for data acquisition and therefore to calibrate the sensor. On the other hand, most of the sensor systems are black box (Figure 5). This means that most of the manufacturers of sensor systems does not commercialize sensors that can be re-calibrated according to the requirements of the user. Sensor systems are intended to be ready-to-use air quality monitors. When purchased by the end-user, a sensor system should estimate the concentration of pollutants with a close agreement to the traditional reference systems used to monitor air quality.

Figure 5. Prices of SS grouped by model. (Numbers in bold indicates the number of pollutant measured by each sensor. x-axis uses logarithmic scale). Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ & ‘updated’ and ‘non-living’ sensors are indicated in black and red color, respectively. NC indicates non commercially available sensor.

In Figure 6 We have shortlisted sensor systems according to their level of agreement to reference systems. For this purpose, we considered metrics obtained from 1 hour and 24 hour averaged data of sensor systems with R2 > 0.85 and 0.8 < slope < 1.2.

Among open source sensor systems we could identify the AirSensEUR (v.2) by LiberaIntentio and the AIRQuino by the CNR for the detection of NO2, CO, O3, NO and PM, respectively. The remaining shortlisted sensor systems were identified as black box. Table 4 and Table S7 report the mean value of R2 and of the slope for the sensors systems shortlisted in Figure 15 for 1 hour and 24 hour averaged data. As we can see, the AirSensEUR (v.2) resulted in a mean R2 value of ~ 0.90 and a slope of ~ 0.94 while the AIRQuino resulted in a mean R2 value of ~ 0.91 and a slope of ~ 0.97. We need to point out that, at the date, the AIRQuino can measures up to five pollutants (PM2.5, PM10, NO2, O3, CO and NO, CO2 and VOCs), however, only data fror PM were available a the time of this review. On the other hand, the AirSensEUR (v2) is a complete sensor system that can also measure particulate matter beside gaseous pollutants including “CO2 and Rn (radon)”. This sensor system is already operative and has undergone multiple calibrations and field tests where measurements of gaseous pollutants showing good agreement with reference measurements.

Figure 6. Price of low-cost sensor systems associated to measuremnts performed at different averaging times. Numbers in bold indicate the number of pollutant measured by open source (blue) and black box (black) sensors. Only records with R2 > 0.85 and 0.8 < slope < 1.2 are shown. Names of ‘living’ & ‘updated’ and ‘non-living’ sensors are indicated in black and red color, respectively. NC indicates non commercially available sensor.

As shown in Table 4, the price of sensor systems ranged from few hundreds of euros to about 9000 euros. We have investigated the possibility of having a relationship between the performance of the sensor (here expressed as R2) and the selling price of the sensor. For this purpose we have compared the mean R2 of all sensor models against their price.

Table 4. Shortlist of sensor systems showing good agreement with reference systems (R2 > 0.85; 0.8 < slope < 1.2) for 1 hour time averaged data.

| model | pollutant | mean R2 | mean slope | mean intercept | open/close | living | commercial | price (EUR) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AirNut | PM2.5 | 0.86 | 0.88 | 8.6 | black box | Y | commercial | 132 |
| PA-I | PM1 | 0.95 | 0.92 | 0.52 | black box | N | commercial | 132 |
| PA-II | PM1 | 0.99 | 0.82 | 1.8 | black box | Y | commercial | 176 |
| Egg (2018) | PM1 | 0.87 | 0.85 | 0.095 | black box | Y | commercial | 219 |
| PATS+ | PM2.5 | 0.96 | 0.92 | 0.05 | black box | Y | commercial | 440 |
| S-500 | NO2, O3 | 0.87 | 1 | 0.27 | black box | Y | commercial | 440 |
| CairClip O3/NO2 | O3 | 0.88 | 0.88 | 12 | black box | Y | commercial | 600 |
| Portable AS-LUNG | PM1 | 0.89 | 0.87 | 1 | black box | Y | non commercial | 880 |
| AirSensEUR (v.2) | NO2, O3, CO, NO | 0.91 | 0.98 | 5.7 | open source | Y | commercial | 1600 |
| Met One (NM) | PM2.5 | 0.86 | 1.1 | 2.8 | black box | Y | commercial | 1672 |
| Air Quality Station | PM1 | 0.88 | 0.9 | 0.85 | black box | Y | non commercial | 1760 |
| AQY v0.5 | PM2.5 | 0.87 | 0.97 | 4 | black box | updated | commercial | 2640 |
| Vaisala AQT410 v.1.15 | CO | 0.87 | 0.97 | 0.23 | black box | Y | commercial | 3256 |
| 2B Tech. (POM) | O3 | 1 | 1 | 0.74 | black box | Y | commercial | 3960 |
| AQMesh v.3.0 | NO | 0.87 | 0.88 | 0.76 | black box | N | commercial | 8800 |

In Figure 7 we reported the relation between the mean R2 and the selling price of OEM/SS for field tests comparisons of sensors against reference systems using 1 hour averaged data. As shown in Figure 7, we did not find a significative relation between the commercial price of OEM sensors and the value of R2 . On the other hand, we could observe a slight increase of the price of sensor systems together with R2. The regression equations indicated in Figure 7 have been calculated only considering “living” (or active) OEM/sensor when compared to reference measurements during field tests. For 24 hour averaged data from both OEMs and SS we did not show any relationship between mean R2 and selling price.

Figure 7. Relation between prices of OEMs/Sensor Systems (SS) and R2 for field test only. Logarithmic scale has been set for both axis. Open source and black box models are indicated with open and full circles, respectively. Names of ‘living’ and ‘non-living’ sensors are indicated in black and blue color, respectively. R2 refers to data averaged over 1 hour. Grey shade in the fit plots indicate a pointwise 95% confidence interval on the fitted values.

As shown in Figure 7, most of the reviewed sensor systems are “black box” systems. This means that the end-user cannot perform any further correction or re-calibration on the sensor system itself. The relationship observed between the coefficient of determination R2 and the price of the sensor system, it is an evidence that complexity of building a detector for air quality is somewhat linked to the choice of materials, multi-functionality and time-spent to develop a reliable sensor system.

In order to target sensor systems in closer agreement and accuracy with reference systems, we displayed the distribution of SS models with mean R2 > 0.75 and 0.5 < mean slope < 1.2. Figure 8 indicates the 2B Tech. (POM) by 2B Technologies, the AirSensEUR (v.2) by Liberantentio, the S-500 by Aeroqual, the Egg (v.2) by Air Quality Egg, the AQT410 v.1.15 by Vaisala and the AirVeraCity as the one having mean R2 –> 1 and $ mean slope = 1. These sensor systems give indicative measurements of air pollutants when comparared with the traditional reference monitoring systems over 1 hour averaging time. On the other hand, other sensors such as the PA-II by Purple Air, the AirNut, the AQMesh v.3.0 by AQMesh and, the AQY v0.5 by Aeroqual showed good agreement with reference systems but lower accuracy.

Figure 8. Correspondence between R2 and slope for sensor systems (SS) for 1 hour averaging time. Only sensor models with mean R2 > 0.75 and 0.5 < mean slope < 1.2 are shown. Names of ‘living’ and ‘non-living’ sensors are indicated in black and blue color, respectively.

## 7. Conclusions

According to the European Air Quality Directive2, a sensor system can be considered “Equivalent” when it meets the Data Quality Objectives (DQOs) set for data capture and uncertainty3,4 .In order for sensor system measurement to be incorporated into the legal framework set by the Air Quality Directive in Europe, they shall satisfy one of the data quality objectives (DQOs) of the Directive. DQOs, defined as the maximum allowed relative uncertainty, are defined either for reference and indicative measurements or for objective estimations. For inorganic gaseous pollutants, they correspond to 15, 25 to 30 and 75 %, respectively. Although, the objective of sensor systems is to provide the most accurate air pollution measurements, it is most likely that the DQO for reference measurements is out reach while it is believed that by improving the sensor calibration procedures the DQO of “Indicative Measurements” could be met at fixed monitoring sites39.

## Supplementary material

Table S1. Number of analyzed records and sensor models by averaging time.

| Averaging time | n. records | n. OEMs & SS |
| --- | --- | --- |
| 1 hour | 610 | 86 |
| 5 min | 253 | 40 |
| 24 hour | 248 | 42 |
| 1 min | 214 | 33 |

Table S2. Metrics used for comparing sensor data, Mi, and reference measurements, RMi.

| Comparison metrics | Short name | Mathematical formulas | Characteristics |
| --- | --- | --- | --- |
| Coefficient of determination | R2 | R2=1−SSRESSSTOT , where SS_RES is the sum of squares of residuals and SS_TOT is the total sum of squares | R2 measures the strentgth of relationship between Mi and RMi, of the percentage of total variance that is explained by a linear relationship |
| Slope of linear relation ship | slope |  |  |
| Intercept of the linear relationship Mi = Slope RMI + Offset | Offset |  |  |
| Root Mean Square Error | RMSE | 1ni=1nMi−RMi2 | indicates the magnitude of the error and retains the variable’s unit; is sensitive to extreme values and to outliers; tends to vary as a function of the standard deviation of the RM |
| Measurement uncertainty | U | U=kUc |  |
| Correlation Coefficient | r | 1ni=1nMi−M)(Mi−RM1ni=1nMi−M21ni=1nRMi−RM | measures the strength and the direction of a linear relationship between two variables, and receives a value between -1 and 1; is independent of the difference in the variance (var) of M and RM, thus if r=1 and var(M)<var(RM), then variance correction may be required |

Table S3. Model of OEMs by pollutant, type, openness and price.

|  | model | pollutant | type | reference | open/close | living | price |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 9 | CO-B4 | CO | electrochemical | Wei[78] | open source | Y | 50 |
| 8 | CO 3E300 | CO | electrochemical | Gerboles[28] | open source | Y | 100 |
| 10 | DataRAM pDR-1200 | PM2.5 | nephelometer | Chakrabarti[19] | black box | N |  |
| 11 | DiscMini | PM | OPC | Viana[74] | open source | Y | 11000 |
| 12 | DN7C3CA006 | PM2.5 | nephelometer | Sousan[59] | open source | Y | 10 |
| 7 | DSM501A | PM2.5 | nephelometer | Wang[75], Alvarado[2] | open source | Y | 15 |
| 13 | FIS SP-61 | O3 | MOs | Spinelle[62] | open source | Y | 50 |
| 14 | GP2Y1010AU0F | PM2.5, PM10 | nephelometer | Olivares[54], Manikonda[47], Sousan[59], Alvarado[2], Wang[75] | open source | Y | 10 |
| 15 | MiCS-2710 | NO2 | MOs | Spinelle[63], Williams[81] | open source | N | 7 |
| 16 | MICS-4514 | CO, NO2 | MOs | Spinelle[64], Spinelle[63] | open source | Y | 20 |
| 6 | NO-3E100 | NO | electrochemical | Spinelle[64], Gerboles[28] | open source | Y | 120 |
| 17 | NO-B4 | NO | electrochemical | Wei[78] | open source | Y | 50 |
| 3 | NO2-3E50 | NO2 | electrochemical | Spinelle[63], Spinelle[61], Gerboles[28] | open source | Y | 100 |
| 2 | NO2-A1 | NO2 | electrochemical | Williams[81] | black box | Y | 50 |
| 18 | NO2-B4 | NO2 | electrochemical | Spinelle[61], Spinelle[63] | open source | N | 50 |
| 19 | NO2-B42F | NO2 | electrochemical | Wei[78] | open source | N | 50 |
| 20 | NO2-B43F | NO2 | electrochemical | Sun[70] | open source | Y | 50 |
| 4 | O3-3E1F | O3 | electrochemical | Spinelle[61], Spinelle[63], Gerboles[28] | open source | Y | 500 |
| 21 | O3-B4 | O3 | electrochemical | Spinelle[61], Spinelle[63], Wei[78] | open source | N | 50 |
| 1 | OPC-N2 | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3], Mukherjee[52], Sousan[60], Feinberg[26], Crilley[21], Di-Antonio[24], Badura[5], LCSQA[42] | black box, open source | N | 310 |
| 27 | OPC-N3 | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 338 |
| 22 | PMS1003 | PM2.5 | OPC | Kelly[37] | black box | Y | 20 |
| 23 | PMS3003 | PM2.5 | OPC | Zheng[83], Kelly[37] | black box | Y | 30 |
| 24 | PMS5003 | PM2.5 | OPC | Laquai[41] | black box | Y | 15 |
| 25 | PMS7003 | PM2.5 | OPC | Badura[5] | black box | Y | 20 |
| 26 | PPD42NS | PM2.5, PM3, PM2 | nephelometer | Wang[75], Holstius[31], Gao[27], Kelly[37], Austin[4] | open source | Y | 15 |
| 28 | SDS011 | PM2.5, PM10 | OPC | Budde[11], Badura[5], Liu[46], Laquai[41] | black box | Y | 30 |
| 29 | SM50 | O3 | MOs | Feinberg[26] | open source | Y | 500 |
| 5 | TGS-5042 | CO | MOs | Spinelle[64] | open source | Y | 40 |
| 30 | TZOA-PM Research Sensors | PM | nephelometer | Feinberg[26] | open source | Y | 90 |
| 31 | ZH03A | PM2.5 | OPC | Badura[5] | black box | Y | 20 |

Table S4. Models of Sensor Systems by pollutant, type, openness and price.

|  | model | pollutant | type | reference | open/close | living | price |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2 | 2B Tech. (POM) | O3 | UV | AQ-SPEC[3] | black box | Y | 4500 |
| 10 | Aeroqual-SM50 | O3 | MOs | Jiao[33] | black box | Y | 2000 |
| 31 | AGT ATS-35 NO2 | NO2 | MOs | Williams[81] | black box | N |  |
| 40 | Air Quality Station | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 2000 |
| 29 | AirAssure | PM2.5 | nephelometer | Feinberg[26], Manikonda[47], AQ-SPEC[3] | black box | Y | 1500 |
| 9 | AirBeam | PM2.5 | OPC, nephelometer | Mukherjee[52], Feinberg[26], Borghi[9], Jiao[33], AQ-SPEC[3], LCSQA[42] | black box | Y | 200 |
| 22 | AirCube | NO2, O3, NO | electrochemical | Mueller[51], Bigi[8] | black box | Y | 3538 |
| 74 | AirMatrix | PM1, PM10, PM2.5 | nephelometer | LCSQA[42] | black box | Y | 60 |
| 52 | AirNut | PM2.5 | OPC | AQ-SPEC[3] | black box | Y | 150 |
| 33 | AIRQino | PM2.5, PM10 | OPC | Cavaliere[13] | black box | Y | 1000 |
| 20 | AirSensEUR (v.1) | CO, NO, NO2, O3 | electrochemical | Wastine[76], Wastine[77], LCSQA[42] | open source, black box | Y | 1600 |
| 26 | AirSensEUR (v.2) | CO, NO, NO2, O3 | electrochemical | Karagulian[36] | open source | Y | 1600 |
| 35 | AirSensorBox | NO2, CO, O3, PM10 | electrochemical, MOs, nephelometer | Borrego[10] | black box | Y | 280 |
| 66 | AirThinx | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 1000 |
| 23 | AirVeraCity | CO, NO2, O3 | electrochemical, MOs | Marjovi[48] | black box | Y | 10000 |
| 55 | AirVisual Pro | PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 270 |
| 11 | AQMesh v.3.0 | CO, NO | electrochemical | Jiao[33] | black box | N | 10000 |
| 3 | AQMesh v.4.0 | CO, NO2, NO, O3, PM1, PM10, PM2.5 | electrochemical, OPC | AQ-SPEC[3], Cordero[20], Castell[12], Borrego[10], LCSQA[42] | black box | Y | 10000 |
| 57 | AQT-410 v.1.11 | O3 | electrochemical | AQ-SPEC[3] | black box | Y | 3700 |
| 76 | AQT-420 | NO2, O3, PM10, PM2.5 | electrochemical, OPC | LCSQA[42] | black box | Y | 5000 |
| 67 | AQY v0.5 | PM2.5, NO2, O3 | OPC, electrochemical, MOs | AQ-SPEC[3] | black box | updated | 3000 |
| 41 | ARISense | NO2, CO, NO, O3 | electrochemical | Cross[22] | black box | Y |  |
| 79 | Atmotrack | PM1, PM10, PM2.5 | OPC | LCSQA[42] | black box | Y | 2500 |
| 46 | BAIRS | PM2.5-0.5 | OPC | Northcross[53] | open source | N | 475 |
| 58 | Cair | PM2.5, PM10-2.5 | OPC | AQ-SPEC[3] | black box | Y | 200 |
| 42 | CairClip NO2-F | NO2 | electrochemical | Spinelle[61], Spinelle[63], Duvall[25], LCSQA[42] | black box | Y | 600 |
| 12 | CairClip O3/NO2 | O3, NO2 | electrochemical | Jiao[33], Spinelle[61], Williams[81], Duvall[25], Feinberg[26] | black box | Y | 600 |
| 43 | CairClip PM2.5 | PM2.5 | OPC | Williams[80] | black box | Y | 1500 |
| 34 | CAM | PM10, PM2.5, NO2, CO, NO | OPC, electrochemical | Borrego[10] | black box | Y |  |
| 44 | CanarIT | PM | OPC | Williams[80] | black box | N | 1500 |
| 70 | Clarity Node | PM2.5 | OPC | AQ-SPEC[3] | black box | Y | 1300 |
| 13 | Dylos DC1100 | PM2.5-0.5 | OPC | Jiao[33], Williams[80], Feinberg[26] | black box, open source | Y | 300 |
| 4 | Dylos DC1100 PRO | PM2.5-0.5, PM10-2.5, PM10 | OPC | AQ-SPEC[3], Jiao[33], Feinberg[26], Manikonda[47] | black box, open source | Y | 300 |
| 45 | Dylos DC1700 | PM2.5-0.5, PM10, PM10-2.5, PM3, PM2, PM2.5 | OPC | Manikonda[47], Sousan[59], Northcross[53], Holstius[31], Steinle[68], Han[30], Jovasevic[34], Dacunto[23] | open source | Y | 475 |
| 81 | e-PM | PM10, PM2.5 | OPC | LCSQA[42] | black box | Y | 2500 |
| 59 | E-Sampler | PM2.5 | OPC | AQ-SPEC[3] | black box | Y | 5500 |
| 47 | ECN_Box | PM10, PM2.5, NO2, O3 | nephelometer, electrochemical | Borrego[10] | black box | Y | 274 |
| 48 | Eco PM | PM1 | OPC | Williams[80] | black box | N |  |
| 75 | ECOMSMART | NO2, O3, PM1, PM10, PM2.5 | electrochemical, OPC | LCSQA[42] | black box | Y | 4650 |
| 65 | Egg (2018) | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 249 |
| 1 | Egg v.1 | CO, NO2, O3 | MOs | AQ-SPEC[3] | black box | N | 200 |
| 27 | Egg v.2 | CO, NO2, O3 | electrochemical | AQ-SPEC[3] | black box | Y | 240 |
| 28 | Egg v.2 (PM) | PM2.5, PM10 | nephelometer | AQ-SPEC[3] | black box | Y | 280 |
| 5 | ELM | NO2, PM10, O3 | MOs, nephelometer | AQ-SPEC[3], US-EPA[72] | black box | N | 5200 |
| 50 | EMMA | PM2.5, CO, NO2, NO | OPC, electrochemical | Gillooly[29] | black box | Y |  |
| 80 | ES-642 | PM2.5 | OPC | LCSQA[42] | black box | Y | 2600 |
| 37 | Foobot | PM2.5 | OPC | AQ-SPEC[3] | black box | Y | 200 |
| 36 | Hanvon N1 | PM2.5 | nephelometer | AQ-SPEC[3] | black box | Y | 200 |
| 19 | Intel Berkeley Badge | NO2, O3 | electrochemical, MOs | Vaughn[73] | open source | N |  |
| 51 | ISAG | NO2, O3 | MOs | Borrego[10] | black box | N |  |
| 38 | Laser Egg | PM2.5, PM10 | nephelometer | AQ-SPEC[3] | black box | Y | 200 |
| 53 | M-POD | CO, NO2 | MOs | Piedrahita[55] | black box | N |  |
| 17 | MAS | CO, NO2, O3, PM2.5 | electrochemical, UV, OPC | Sun[69] | black box, open source | N, Y | 5500 |
| 54 | Met One - 831 | PM10 | OPC | Williams[80] | black box | Y | 2050 |
| 6 | Met One (NM) | PM2.5 | OPC | AQ-SPEC[3], LCSQA[42] | black box | Y | 1900 |
| 7 | MicroPEM | PM2.5 | OPC | AQ-SPEC[3], Williams[80] | black box | Y | 2000 |
| 56 | NanoEnvi | NO2, O3, CO | electrochemical, MOs | Borrego[10] | black box | Y |  |
| 39 | PA-I | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | N | 150 |
| 73 | PA-I-Indoor | PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 180 |
| 60 | PA-II | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 200 |
| 8 | Partector | PM1, PM2.5 | Electrical | AQ-SPEC[3] | black box | Y | 7000 |
| 68 | PATS+ | PM2.5 | OPC | Pillarisetti[56] | black box | Y | 500 |
| 69 | Platypus NO2 | NO2 | MOs | Williams[81] | black box | Y | 50 |
| 14 | PMS-SYS-1 | PM2.5 | nephelometer | Jiao[33], AQ-SPEC[3], Williams[80], Feinberg[26] | black box | Y | 1000 |
| 49 | Portable AS-LUNG | PM1, PM2.5, PM10 | OPC | AQ-SPEC[3] | black box | Y | 1000 |
| 61 | Pure Morning P3 | PM2.5 | OPC | AQ-SPEC[3] | black box | Y | 170 |
| 21 | RAMP | CO, NO2 | electrochemical | Zimmerman[85] | open source | Y |  |
| 16 | S-500 | O3, NO2 | MOs | AQ-SPEC[3], Lin[45], Vaughn[73] | black box | Y | 500 |
| 24 | SENS-IT | O3, CO, NO2 | MOs | AQ-SPEC[3] | black box | N, Y | 2200 |
| 71 | SidePak AM510 | PM2.5 | nephelometer | Karagulian[35] | open source | N | 3000 |
| 25 | Smart Citizen Kit | CO | MOs | AQ-SPEC[3] | black box | Y | 200 |
| 30 | SNAQ | NO2, CO, NO | electrochemical | Mead[49], Popoola[57] | black box | Y |  |
| 32 | Spec | CO, NO2, O3 | electrochemical | AQ-SPEC[3] | black box | Y | 500 |
| 15 | Speck | PM2.5 | nephelometer | Feinberg[26], US-EPA[72], Williams[80], Manikonda[47], Zikova[84], AQ-SPEC[3] | black box | Y | 150 |
| 72 | UBAS | PM2.5 | nephelometer | Manikonda[47] | black box | N | 100 |
| 62 | uHoo | PM2.5, O3 | nephelometer, MOs | AQ-SPEC[3] | black box | Y | 300 |
| 18 | Urban AirQ | NO2 | electrochemical | Mijling[50] | open source | N |  |
| 63 | Vaisala AQT410 v.1.11 | CO, NO2 | electrochemical | AQ-SPEC[3] | black box | Y | 3700 |
| 64 | Vaisala AQT410 v.1.15 | CO, NO2 | electrochemical | AQ-SPEC[3] | black box | Y | 3700 |
| 77 | Waspmote | NO, NO2, PM1, PM10, PM2.5 | MOs, OPC | LCSQA[42] | open source | Y | 1270 |
| 78 | Watchtower 1 | NO2, PM1, PM10, PM2.5, O3 | electrochemical, OPC | LCSQA[42] | black box | Y | 5000 |

Figure S1. Distribution of R2 for OEMs and sensor systems obtained from the calibration against the reference. Records were averaged over a time-scale of 1 hour. Dashed lines indicate the value of 0.7 and 1.0. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S2. Distribution of R2 for OEMs and ensor systems obtained from the calibration against the reference. Records were averaged over a time-scale of 1 minute. Dashed lines indicate the value of 0.7 and 1.0. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S3. Distribution of R2 from the comparison of sensor systems against reference systems. Records were averaged over a time-scale of 1 minute. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S4. Distribution of R2 from the comparison of sensor systems against reference systems. Records were averaged over a time-scale of 1 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S5. Distribution of R2 from the comparison of OEMs against reference systems. Records were averaged over a time-scale of 1 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S6. Distribution of R2 from the comparison of OEMs against reference systems. Records were averaged over a time-scale of 24 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S7. Distribution of R2 from the comparison of sensor systems against reference systems. Records were averaged over a time-scale of 24 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S8. Distribution of slopes from the comparison of sensors systems against the reference. Only records with R2 > 0.7 and 0.5 < slope < 1.5 are shown. Records were averaged over a time-scale of 1 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S9. Distribution of slopes from the comparison of OEMs against the reference. Only records with R2 > 0.7 and 0.5 < slope < 1.5 are shown. Records were averaged over a time-scale of 1 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S10. Distribution of slopes from the comparison of sensors systems against the reference. Only records with R2 > 0.7 and 0.5 < slope < 1.5 are shown. Records were averaged over a time-scale of 24 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S11. Distribution of slopes from the comparison of OEMs against the reference. Only records with R2 > 0.7 and 0.5 < slope < 1.5 are shown. Records were averaged over a time-scale of 24 hour. Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Figure S12. Mean slope for obtained from the comparison of OEMs and sensor systems against reference measurements.

Figure S13. Prices of OEMs available on the market (Numbers in bold indicates the number of pollutant measured by each sensor. x-axis uses logarithmic scale). Numbers in bold indicate the number of open source (blue) and black box (black) records. Names of ‘living’ and ‘non-living’ sensors are indicated in black and red color, respectively.

Table S7. Shortlist of sensor systems showing good agreement with reference systems (R2 > 0.85; 0.8 < slope < 1.2) for 24 hour time averaged data.

| model | pollutant | mean R2 | mean slope | mean intercept | open/close | living | commercial | price (EUR) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PA-I | PM1 | 0.99 | 0.91 | 0.47 | black box | N | commercial | 132 |
| PA-II | PM1 | 0.99 | 0.83 | 1.8 | black box | Y | commercial | 176 |
| Egg (2018) | PM1 | 0.88 | 0.81 | 0.33 | black box | Y | commercial | 219 |
| Egg v.2 (PM) | PM2.5 | 0.94 | 1 | 3.3 | black box | Y | commercial | 246 |
| AirThinx | PM1 | 0.89 | 0.85 | 1.3 | black box | Y | commercial | 880 |
| Portable AS-LUNG | PM1 | 0.93 | 0.88 | 1.5 | black box | Y | non commercial | 880 |
| AIRQino | PM2.5, PM10 | 0.93 | 1 | 1.1 | black box | Y | non commercial | 1000 |
| Air Quality Station | PM1 | 0.94 | 0.89 | 1.1 | black box | Y | non commercial | 1760 |
| AQY v0.5 | PM2.5 | 0.91 | 0.94 | 4 | black box | updated | commercial | 2640 |
| Vaisala AQT410 v.1.15 | CO | 0.86 | 0.91 | 0.25 | black box | Y | commercial | 3256 |

## Bibliography

1. Aleixandre, M. & Gerboles, M. Review of small commercial sensors for indicative monitoring of ambient gas. Chemical Engineering Transactions 30, 169–174 (2012).

2. Alvarado, M. et al. Towards the Development of a Low Cost Airborne Sensing System to Monitor Dust Particles after Blasting at Open-Pit Mine Sites. Sensors 15, 19667–19687 (2015).

3. AQ-SPEC, District, S. C. A. Q. M. & District, S. C. A. Q. M. Air Quality Sensor Performance Evaluation Reports. (2015).

4. Austin, E., Novosselov, I., Seto, E. & Yost, M. G. Laboratory Evaluation of the Shinyei PPD42NS Low-Cost Particulate Matter Sensor. PLOS ONE 10, e0137789 (2015).

5. Badura, M., Batog, P., Drzeniecka-Osiadacz, A. & Modzel, P. Optical particulate matter sensors in PM 2.5$ measurements in atmospheric air. E3S Web of Conferences 44, 00006 (2018).

6. Barrett, J. P. The Coefficient of Determination—Some Limitations. The American Statistician 28, 19–20 (1974).

7. Bettair. Bettair - Solving Cities Air Pollution bettair cities.

8. Bigi, A., Mueller, M., Grange, S. K., Ghermandi, G. & Hueglin, C. Performance of NO, NO2 low cost sensors and three calibration approaches within a real world application. Atmospheric Measurement Techniques 11, 3717–3735 (2018).

9. Borghi, F. et al. Precision and Accuracy of a Direct-Reading Miniaturized Monitor in PM2.5 Exposure Assessment. Sensors 18, 3089 (2018).

10. Borrego, C. et al. Assessment of air quality microsensors versus reference methods: The EuNetAir joint exercise. Atmospheric Environment 147, 246–263 (2016).

11. Budde, M. et al. Suitability of the Low-Cost SDS011 Particle Sensor for Urban PM-Monitoring. (2018).

12. Castell, N. et al. Can commercial low-cost sensor platforms contribute to air quality monitoring and exposure estimates? Environment International 99, 293–302 (2017).

13. Cavaliere, A. et al. Development of Low-Cost Air Quality Stations for Next Generation Monitoring Networks: Calibration and Validation of PM2.5 and PM10 Sensors. Sensors 18, 2843 (2018).

14. CEN. Ambient air - Standard method for the measurement of the concentration of carbon monoxide by non-dispersive infrared spectroscopy, (EN 14626:2012). European Committee for Standardization. (2012).

15. CEN. Ambient air - Standard method for the measurement of the concentration of nitrogen dioxide and nitrogen monoxide by chemiluminescence (EN 14211:2012). European Committee for Standardization. (2012).

16. CEN. Ambient air - Standard method for the measurement of the concentration of ozone by ultraviolet photometry, (EN 14625:2012). European Committee for Standardization. (2012).

17. CEN. Ambient air - Standard method for the measurement of the concentration of sulphur dioxide by ultraviolet fluorescence, (EN 14212:2012). European Committee for Standardization. (2012).

18. CEN. Ambient air - Standard gravimetric measurement method for the determination of the PM10 or PM2,5 mass concentration of suspended particulate matter (EN 12341:2014). European Committee for Standardization. (2014).

19. Chakrabarti, B., Fine, P. M., Delfino, R. & Sioutas, C. Performance evaluation of the active-flow personal DataRAM PM2.5 mass monitor (Thermo Anderson pDR-1200) designed for continuous personal exposure measurements. Atmospheric Environment 38, 3329–3340 (2004).

20. Cordero, J. M., Borge, R. & Narros, A. Using statistical methods to carry out in field calibrations of low cost air quality sensors. Sensors and Actuators B: Chemical 267, 245–254 (2018).

21. Crilley, L. R. et al. Evaluation of a low-cost optical particle counter (Alphasense OPC-N2) for ambient air monitoring. Atmospheric Measurement Techniques 11, 709–720 (2018).

22. Cross, E. S. et al. Use of electrochemical sensors for measurement of air pollution: Correcting interference response and validating measurements. Atmos. Meas. Tech. 14 (2017).

23. Crunaire, S., Redon, N. & Spinelle, L. 1ER ESSAI NATIONAL D’APTITUDE DES MICRO-CAPTEURS(EAμC) POUR LA SURVEILLANCE DE LA QUALITE DE L’AIR : SYNTHESE DES RESULTATS. 38 (LCSQA, 2018).

24. Dacunto, P. J. et al. Determining PM 2.5$ calibration curves for a low-cost particle monitor: Common indoor residential aerosols. Environmental Science: Processes & Impacts 17, 1959–1966 (2015).

25. Di Antonio, A., Popoola, O. A. M., Ouyang, B., Saffell, J. & Jones, R. L. Developing a Relative Humidity Correction for Low-Cost Sensors Measuring Ambient Particulate Matter. Sensors 18, 2790 (2018).

26. Duvall, R. M. et al. Performance Evaluation and Community Application of Low-Cost Sensors for Ozone and Nitrogen Dioxide. Sensors 16, 1698 (2016).

27. Feinberg, S. et al. Long-term evaluation of air sensor technology under ambient conditions in Denver, Colorado. Atmospheric Measurement Techniques 11, 4605–4615 (2018).

28. Gao, M., Cao, J. & Seto, E. A distributed network of low-cost continuous reading sensors to measure spatiotemporal variations of PM2.5 in Xi’an, China. Environmental Pollution 199, 56–65 (2015).

29. Gerboles, M., Spinelle, L. & Signorini, M. AirSensEUR: An open data/software/hardware multi-sensor platform for air quality monitoring. Part A: Sensor shield. (Publications Office of the European Union, 2015). doi:10.2788/30927

30. Gillooly, S. E. et al. Development of an in-home, real-time air pollutant sensor platform and implications for community use. Environmental Pollution 244, 440–450 (2019).

31. Han, I., Symanski, E. & Stock, T. H. Feasibility of using low-cost portable particle monitors for measurement of fine and coarse particulate matter in urban ambient air. Journal of the Air & Waste Management Association 67, 330–340 (2017).

32. Holstius, D. M., Pillarisetti, A., Smith, K. R. & Seto, E. Field calibrations of a low-cost aerosol sensor at a regulatory monitoring site in California. Atmospheric Measurement Techniques 7, 1121–1131 (2014).

33. iScape. Summary of Air Quality sensors and recommendations for application. iScape project D1.5, February 2017. (2017).

34. Jiao, W. et al. Community Air Sensor Network (CAIRSENSE) project: Evaluation of low-cost sensor performance in a suburban environment in the southeastern United States. Atmospheric Measurement Techniques 9, 5281–5292 (2016).

35. Jovašević-Stojanović, M. et al. On the use of small and cheaper sensors and devices for indicative citizen-based monitoring of respirable particulate matter. Environmental Pollution 206, 696–704 (2015).

36. Karagulian, F. et al. Calibration of AirSensEUR units during a field study in the Netherlands. (European Commission - Joint Research Centre, 2019).

37. Karagulian, F., Belis, C. A., Lagler, F., Barbiere, M. & Gerboles, M. Evaluation of a portable nephelometer against the Tapered Element Oscillating Microbalance method for monitoring PM2.5. Journal of Environmental Monitoring 14, 2145–2153 (2012).

38. Kelly, K. E. et al. Ambient and laboratory evaluation of a low-cost particulate matter sensor. Environmental Pollution 221, 491–500 (2017).

39. KEY-VOCs. Metrology for VOC indicators in air pollution and climate change. Revision of the testing protocol and procedure for calibration based on laboratory validation, final report. ENV56 KEY-VOCs. (2017).

40. Kumar, P. et al. The rise of low-cost sensing for managing air pollution in cities. Environment International 75, 199–205 (2015).

41. Kunak. Wireless Particle counter for Outdoors Kunak AIR P10. Kunak

42. Laquai, B. Particle Distribution Dependent Inaccuracy of the Plantower PMS5003 low-cost PM-sensor. (2017).

43. Lewis, A. C., Schneidemesser, E. von & Peltier, R. Low-cost sensors for the measurement of atmospheric composition: Overview of topic and future applications (World Meteorological Organization) WMO-No. 1215. (2018).

44. Lewis, A. & Edwards, P. Validate personal air-pollution sensors. Nature News 535, 29 (2016).

45. Lin, C. et al. Evaluation and calibration of Aeroqual series 500 portable gas sensors for accurate measurement of ambient ozone and nitrogen dioxide. Atmospheric Environment 100, 111–116 (2015).

46. Liu, H.-Y., Schneider, P., Haugen, R. & Vogt, M. Performance Assessment of a Low-Cost PM2.5 Sensor for a near Four-Month Period in Oslo, Norway. Atmosphere 10, 41 (2019).

47. Manikonda, A., Zíková, N., Hopke, P. K. & Ferro, A. R. Laboratory assessment of low-cost PM monitors. Journal of Aerosol Science 102, 29–40 (2016).

48. Marjovi, A., Arfire, A. & Martinoli, A. Extending Urban Air Quality Maps Beyond the Coverage of a Mobile Sensor Network: Data Sources, Methods, and Performance Evaluation. in Proceedings of the 2017 International Conference on Embedded Wireless Systems and Networks 12–23 (Junction Publishing, 2017).

49. Mead, M. I. et al. The use of electrochemical sensors for monitoring urban air quality in low-cost, high-density networks. Atmospheric Environment 70, 186–203 (2013).

50. Mijling, B., Jiang, Q., Jonge, D. de & Bocconi, S. Practical field calibration of electrochemical NO2 sensors for urban air quality applications. Atmos. Meas. Tech. Discuss. 2017, 1–25 (2017).

51. Mueller, M., Meyer, J. & Hueglin, C. Design of an ozone and nitrogen dioxide sensor unit and its long-term operation within a sensor network in the city of Zurich. Atmos. Meas. Tech. 17 (2017).

52. Mukherjee, A., Stanton, L. G., Graham, A. R. & Roberts, P. T. Assessing the Utility of Low-Cost Particulate Matter Sensors over a 12-Week Period in the Cuyama Valley of California. Sensors 17, 1805 (2017).

53. Northcross, A. L. et al. A low-cost particle counter as a realtime fine-particle mass monitor. Environmental Science: Processes & Impacts 15, 433–439 (2013).

54. Olivares, G. & Edwards, S. The Outdoor Dust Information Node (ODIN) – development and performance assessment of a low cost ambient dust sensor. Atmospheric Measurement Techniques Discussions 8, 7511–7533 (2015).

55. Piedrahita, R. et al. The next generation of low-cost personal air quality sensors for quantitative exposure monitoring. Atmospheric Measurement Techniques 7, 3325–3336 (2014).

56. Pillarisetti, A. et al. Small, Smart, Fast, and Cheap: Microchip-Based Sensors to Estimate Air Pollution Exposures in Rural Households. Sensors (Basel, Switzerland) 17, (2017).

57. Popoola, O. A. M., Stewart, G. B., Mead, M. I. & Jones, R. L. Development of a baseline-temperature correction methodology for electrochemical sensors and its implications for long-term stability. Atmospheric Environment 147, 330–343 (2016).

58. Snyder, E. et al. The changing paradigm of air pollution monitoring. Environmental Science and Technology 47, 11369–11377 (2013).

59. Sousan, S., Koehler, K., Hallett, L. & Peters, T. M. Evaluation of the Alphasense optical particle counter (OPC-N2) and the Grimm portable aerosol spectrometer (PAS-1.108). Aerosol Science and Technology 50, 1352–1365 (2016).

60. Sousan, S. et al. Inter-comparison of low-cost sensors for measuring the mass concentration of occupational aerosols. Aerosol Science and Technology 50, 462–473 (2016).

61. Spinelle, L., Gerboles, M., Aleixandre, M. & Bonavitacola, F. Evaluation of metal oxides sensors for the monitoring of o3 in ambient air at ppb level. Chemical Engineering Transactions 319–324 (2016). doi:10.3303/CET1654054

62. Spinelle, L., Gerboles, M., Kotsev, A. & Signorini, M. Evaluation of low-cost sensors for air pollution monitoring: Effect of gaseous interfering compounds and meteorological conditions. (Publications Office of the European Union, 2017). doi:10.2760/548327

63. Spinelle, L., Gerboles, M., Villani, M. G., Aleixandre, M. & Bonavitacola, F. Field calibration of a cluster of low-cost available sensors for air quality monitoring. Part A: Ozone and nitrogen dioxide. Sensors and Actuators B: Chemical 215, 249–257 (2015).

64. Spinelle, L., Gerboles, M., Villani, M. G., Aleixandre, M. & Bonavitacola. Field calibration of a cluster of low-cost available sensors for air quality monitoring. Part. B: NO, CO and CO2. Sensors and Actuators B: Chemical in preparation,

65. Spinelle, L., Gerboles, M. & Aleixandre, M. Performance Evaluation of Amperometric Sensors for the Monitoring of O 3 and NO 2 in Ambient Air at ppb Level. Procedia Engineering 120, 480–483 (2015).

66. Spinelle, L., Gerboles, M. & Aleixandre, M. Report of laboratory and in-situ validation of micro-sensor for monitoring ambient air pollution - O12: CairClipO3/NO2 of CAIRPOL (F). (Publications Office of the European Union, 2013).

67. Spinelle, L., Gerboles, M. & Aleixandre, M. Report of laboratory and in-situ validation of micro-sensor for monitoring ambient air pollution - NO9: CairClipNO2 of CAIRPOL (F). (Publications Office of the European Union, 2013).

68. Steinle, S. et al. Personal exposure monitoring of PM 2.5 in indoor and outdoor microenvironments. Science of The Total Environment 508, 383–394 (2015).

69. Sun, L. et al. Development and Application of a Next Generation Air Sensor Network for the Hong Kong Marathon 2015 Air Quality Monitoring. Sensors (Basel, Switzerland) 16, (2016).

70. Sun, L., Westerdahl, D. & Ning, Z. Development and Evaluation of A Novel and Cost-Effective Approach for Low-Cost NO2 Sensor Drift Correction. Sensors 17, 1916 (2017).

71. Index, T. W. A. Q. Sensing the Air Quality: Research on Air Quality Sensors. http://aqicn.org/sensor/. (2019).

72. Agency, U. S. E. P. Evaluation of Elm and Speck Sensors. EPA/600/R-15/314 November 2015 www.epa.gov/ord, (2015).

73. Vaughn, D. L., Dye, T. S., Roberts, L., Ray, A. E. & DeWinter, J. L. Characterization of Low-Cost NO2 Sensors. 10 (U.S. Environmental Protection Agency Research Triangle Park, North Carolina, 2010).

74. Viana, M. et al. Field comparison of portable and stationary instruments for outdoor urban air exposure assessments. Atmospheric Environment 123, Part A, 220–228 (2015).

75. Wang, Y. et al. Laboratory Evaluation and Calibration of Three Low-Cost Particle Sensors for Particulate Matter Measurement. Aerosol Science and Technology 49, 1063–1077 (2015).

76. Wastine, B. AirSensEur: Point sur les expérimentations menées depuis 2018. (Atmo-Normadie, 2019).

77. Wastine, B. Essai d’Aptitude AirSensEUR du 12-janv au 22-fev 2018 réalisé par Atmo Normandie pour l’exercise d’intercomparaison n°1 du LCSQA. 28 (Atmo Normandie, 2019).

78. Wei, P. et al. Impact Analysis of Temperature and Humidity Conditions on Electrochemical Sensor Response in Ambient Air Quality Monitoring. Sensors 18, 59 (2018).

79. White, R. et al. Sensors and ’apps’ for community-based: Atmospheric monitoring. EM: Air and Waste Management Association’s Magazine for Environmental Managers 36–40 (2012).

80. Williams, R. et al. Air Sensor Guidebook, United States Environmental Protection Agency (US-EPA). (2014).

81. Williams, R., Kaufman, A., Hanley, T., Rice, J. & Garvey, S. Evaluation of Field-deployed Low Cost PM Sensors. (2014).

82. Williams, R. et al. Sensor Evaluation Report. (2014).

83. Zheng, T. et al. Field evaluation of low-cost particulate matter sensors in high and low concentration environments. Atmospheric Measurement Techniques Discussions 1–40 (2018). doi:10.5194/amt-2018-111

84. Zhou, X., Lee, S., Xu, Z. & Yoon, J. Recent Progress on the Development of Chemosensors for Gases. Chemical Reviews 115, 7944–8000 (2015).

85. Zikova, N. et al. Estimating Hourly Concentrations of PM2.5 across a Metropolitan Area Using Low-Cost Particle Monitors. Sensors (Basel, Switzerland) 17, (2017).

86. Zimmerman, N. et al. A machine learning calibration model using random forests to improve sensor performance for lower-cost air quality monitoring. Atmos. Meas. Tech. 11, 291–313 (2018).