## Slide 1: Review of Low-Cost Sensors for Air Quality (Administrative Arrangement AQSens Commission internal project between DG ENV and DG JRC)
Federico Karagulian1, Michel Gerboles1, Annette Borowiak1
Workshop:
The use of sensors and their possible role in air quality assessment
Brussels, 22 January 2019
1 EC - JRC

## Slide 2
Purpose of the review work:
selecting the best sensor system candidate
for the AQSens project
Gathering quantitative information about sensor systems including:
Agreement between Low-Cost Sensors (LCSs) and reference measurements
Availability of raw data, transparent data treatment and, possibility of a-posteriori calibration
Capability to measure multiple pollutants
Value for money of sensor systems
Interoperability of data according to the INSPIRE directive
Automatic data-transfer and web visualization of sensor data

## Slide 3
Review was focused on sensors measuring PM, O3, NO2, NO and CO concentrations
Review about the use of Low Cost Sensors (LCSs) for Air Quality has been carried out using:
  Scopus database (collected into a Zotero DB and AirMonTech web site (JRC http://db-airmontech.jrc.ec.europa.eu/search.aspx, supplied with Scopus and ResearchGate websites)
  Reports provided by AQ-SPEC, US-EPA, EC, and relevant testing laboratories
Records were created for each laboratory and field test about calibration and validation of sensors
Methodology
e-reporting
Google-sheet (pwd. protected)
Comparison of sensor
based on their metrics

## Slide 4
OEM sensor
Chemical cell or physical unit that produces an analytically useful signal by detecting or measuring the analyte.
Sensor System (SS) or sensor node
Integrated set of hardware that uses one or more sensors to detect and/or measure a chemical concentration or quantity that is able to supply real time measurements.
Main classification of LCS:
“open source”
Data acquisition and data processing can be operated by an “open source” software tuned according to different calibration parameters and environmental conditions.
“black box”
Impossibility to change any parameters of the sensor set by the manufacturer and data treatment.
n. LSCs | 
OEM | 39
SS | 70
open source sensors | 40
black box sensors | 73
commercial
Possibility to be purchased by any end-user
non-commercial
Only used for research studies

## Slide 5
metrics |  | n. records | 
 |  | Field tests | Laboratory tests
 |  | 1144 | 137
R2 | calibration | 213 | 64
 | comparison | 1022 | 72
Slope of reg. line |  | 922 | 55
RMSE |  | 159 | 5
Uncertainty (U) |  | 42 | 30
Price |  | 114 | 
Evaluation criteria
Collection of relevant metric parameters to check performance of sensor
Comparison of same metrics from different sensors
Records on calibration and validation of sensors were created for each laboratory and field test
Limit of R2: measure of strength of associations and not accuracy of measurement. R2 change with the range of measurements  use together with slope
Strength of R²: it is always available

## Slide 6
Pollutant | Type | n. records | n. references
CO | electrochemical | 64 | 19
CO | MOs | 23 | 3
NO | electrochemical | 47 | 15
NO2 | electrochemical | 134 | 29
NO2 | MOs | 34 | 8
O3 | electrochemical | 61 | 17
O3 | MOs | 55 | 7
O3 | UV | 10 | 2
PM2.5 | Electrical | 6 | 1
PM2.5 | nephelometer | 216 | 22
PM2.5 | OPC | 346 | 25
PM1 | Electrical | 6 | 1
PM1 | OPC | 92 | 4
PM10 | nephelometer | 32 | 3
PM10 | OPC | 155 | 13
Every single reference might contain multiple tests (records)
Types of sensors

## Slide 7
Large availability of active (“living”) & commercial sensors
All OEMs and most SS are commercially available. SS that are not commercial are for publication purposes
Most of commercial SS are “black box” systems
Type of LCS and number of tests

## Slide 8
Shortlist of sensor systems (SS) showing R2 > 0.8 and 0.8 < slope < 1.2 when compared against reference measurements (names of 'living' and 'non-living' sensors are indicated in black and red color, respectively)
Price of sensor systems with best R² and slope recorded
Number of pollutant
measured by each sensor
With R²> 0.85 and 0.8 <slope < 1.2
non commercially
available
* Black number -> black box
* Blue number -> open source
*Vailsala 5000E , South cost Science, KUNAKAIR and Bettair were not included in the comparison because of unavailability of
Peer reviewed works

## Slide 9
* full circles -> “black box”
* Open circles -> “open source”
* Living/non-living sensors are indicated
with black and blue colors, respectively
Price for value (only field tests)
Non-relevant dependency of R2 with the price was observed for OEMs
Validation of sensors systems against reference measurement showed a slight dependency of R2 with the price of the sensor.
R² is not accuracy

## Slide 10
Towards the best performing sensor
*names of 'living'/'non-living' sensors are indicated in black/blue color, respectively
Targeting sensor showing mean R2 ~ 1 and mean slope ~ 1 when compared with reference systems
Not all sensor showing mean R2 ~ 1 had necessarily mean slope ~ 1

## Slide 11
model | pollutant | mean R2 | slope | open/close | living | commercial | price (EUR)
AirNut | PM2.5 | 0.8618 | 0.8838 | black box | Y | commercial | 132
PA-I | PM1 | 0.9464 | 0.9234 | black box | N | commercial | 132
PA-II | PM1 | 0.9874 | 0.823 | black box | Y | commercial | 176
Egg (2018) | PM1 | 0.8736 | 0.8484 | black box | Y | commercial | 219
PATS+ | PM2.5 | 0.96 | 0.92 | black box | Y | commercial | 440
S-500 | NO2, O3 | 0.88 | 1.03 | black box | Y | commercial | 440
Portable AS-LUNG | PM1 | 0.8858 | 0.8678 | black box | Y | non commercial | 880
AirSensEUR (v.2) | NO2, O3, CO, NO | 0.8938 | 0.9425 | open source | Y | commercial | 1000
Air Quality Station | PM1 | 0.88 | 0.896 | black box | Y | non commercial | 1760
AQY v0.5 | PM2.5 | 0.8654 | 0.9673 | black box | updated | commercial | 2640
Vaisala AQT410 v.1.15 | CO | 0.8734 | 0.97 | black box | Y | commercial | 3256
2B Tech. (POM) | O3 | 0.9972 | 1.007 | black box | Y | commercial | 3960
AQMesh v.3.0 | NO | 0.87 | 0.883 | black box | N | commercial | 8800
model | pollutant | mean R2 | mean slope | open/close | living | commercial | price (EUR)
PA-I | PM1 | 0.9867 | 0.9086 | black box | N | commercial | 132
PA-II | PM1 | 0.9937 | 0.8252 | black box | Y | commercial | 176
Egg (2018) | PM1 | 0.8776 | 0.8145 | black box | Y | commercial | 219
Egg v.2 (PM) | PM2.5 | 0.9375 | 1.001 | black box | Y | commercial | 246
AirThinx | PM1 | 0.8858 | 0.8483 | black box | Y | commercial | 880
Portable AS-LUNG | PM1 | 0.9347 | 0.876 | black box | Y | non commercial | 880
AIRQino | PM2.5, PM10 | 0.909 | 0.972 | open source | Y | non commercial | 1000
Air Quality Station | PM1 | 0.94 | 0.89 | black box | Y | non commercial | 1760
AQY v0.5 | PM2.5 | 0.9092 | 0.9393 | black box | updated | commercial | 2640
Vaisala AQT410 v.1.15 | CO | 0.8612 | 0.9105 | black box | Y | commercial | 3256
We have shortlisted sensor systems showing good agreement with reference systems (R2 > 0.85; 0.8 < slope < 1.2)
Only few sensors are “open source” and are able to measure multiple pollutants
1 hour time average
24 hour time average

## Slide 12: Conclusions
Availability of an open-access repository reporting
relevant metric for OEMs and Sensor Systems.
Possibility to share the above-mentioned repository
among Entities performing laboratory and field tests
according to a well-defined protocol to calculate
relevant metrics.
Necessity to target the best sensor system by its
openness, agreement with reference systems,
number of measured pollutants and price.
The information reported in this work about price of sensors, is merely indicative and might change with time. The authors do not take any responsibility should prices and performances of sensors should vary.

## Slide 13: Thanks for your attention!